## MetBAM Checklist South
KCBX TERMINALS CO SOUTH |  | SITE VISIT DATE
| 2017-03-03 | 2017-03-08 | 2017-03-16 | 2017-03-20 | 2017-03-24 | 2017-03-27 | 2017-03-31
Meteorological Station
Arrival time: | 9:40 | 10:15 | 8:40 | 8:55 | 7:00 | 9:00 | 8:20
Is the temperature aspirator fan running? | bar: 2017-03-03 to 2017-03-31
Are all signal cables secured and intact? | bar: 2017-03-03 to 2017-03-31
Are wind vane/ anemometer cups intact and operational? | bar: 2017-03-03 to 2017-03-31
Is datalogger date/time correct within 60 sec? | bar: 2017-03-03 to 2017-03-31
Are conditions shown on logger consistent with ambient observations? | bar: 2017-03-03 to 2017-03-31
BAM-1020 PM10 Monitors
Arrival time: | 9:40 | 10:15 | 8:40 | 8:55 | 7:00 | 9:00 | 8:20
Any status error flags showing? | bar: 2017-03-03 to 2017-03-31
Are all monitor dates/times correct within 60 sec? | bar: 2017-03-03 to 2017-03-31
Are all filter transport systems operational? | bar: 2017-03-03 to 2017-03-31
Are all tape reel cover knobs secure? | bar: 2017-03-03 to 2017-03-31
Are hourly data being collected? | bar: 2017-03-03 to 2017-03-31
Is sampler inlet and downtube clean? | bar: 2017-03-03 to 2017-03-31
Is water collection bottle clean? | bar: 2017-03-03 to 2017-03-31
Are O-rings in good condition? | bar: 2017-03-03 to 2017-03-31
Comments section
3/1, ST-SW data logger failed due to rain getting into the enclosure. BAM data OK.
3/3, Installed replacement data logger at ST-SW. Seems to be working fine.
3/8, Cleaned BAM PM10 heads.
3/16, Performed BAM QC checks.
3/16, Changed time on the ST-N BAM from 09:06:55 to 09:08 CST.
3/16, Changed time on the ST-NE BAM from 09:28:05 to 09:27 CST.
3/20, Cleaned enclosure AC filters.
3/20, Changed time on the ST-SW BAM from 09:43:05 to 09:42 CST.
3/24, Audits by J. Sorkness.
3/27, Changed BAM filter tapes.
3/31, Performed BAM QC checks.
Site ID List: | 2017-03-08: MAINTENANCE LIST FOR METEOROLOGICAL STATION
ST-NW = Northwest BAM site | 2017-03-08: Each Visit:; 2017-03-16: Inspect sensors and cables for signs of deterioration or damage and replace as needed
ST-N = North BAM site | 2017-03-08: Monthly:; 2017-03-16: Inspect temperature aspirator and remove dust accumulation | Rev 3 (5/2014)
ST-NE = Northeast BAM site (Partisol samplers) | 2017-03-16: Inspect tower for plumb, and verify that all fasteners are intact and tight
ST-M = Meteorological Monitoring Site | 2017-03-08: Biannually:; 2017-03-16: Perform calibration check on sensors
ST-CE = Center East BAM site | 2017-03-16: Check wind sensor bearings
ST-SW = Southwest BAM site | 2017-03-08: Annually:; 2017-03-16: Replace wind sensor bearings
MAINTENANCE LIST FOR BAM-1020
|  | Each Visit: | Inspect filter tape and replace when needed (~2 months/roll; run SELF-TEST function in the TAPE menu).
Inspect filter tape for pinholes and irregular PM deposits
|  | Monthly: | Perform flow rate and leak check
Clean BAM nozzle and vane
Clean inlet hardware, capstan shaft and pinch roller
|  | Quarterly: | Verify BAM settings file
Flow system calibration
Disassemble and clean inlet
|  | Biannually: | Check pump muffler and tubing and replace as needed
Test filter RH and Temperature sensors
Test smart heater function
|  | Annually: | Perform zero background test
Clean internal debris filter
Remove and check membrane span foil
Beta detector count rate and dark count test
Rebuild vacuum pump
Replace nozzle o-ring