## MetBAM Checklist South
KCBX TERMINALS CO SOUTH |  | SITE VISIT DATE
| 2018-02-02 | 2018-02-06 | 2018-02-13 | 2018-02-16 | 2018-02-20 | 2018-02-23 | 2018-02-26
Meteorological Station
Arrival time: | 10:40 | 11:20 | 10:45 | 10:05 | 11:00 | 7:55 | 10:00
Is the temperature aspirator fan running? | bar: 2018-02-02 to 2018-02-26
Are all signal cables secured and intact? | bar: 2018-02-02 to 2018-02-26
Are wind vane/ anemometer cups intact and operational? | bar: 2018-02-02 to 2018-02-26
Is datalogger date/time correct within 60 sec? | bar: 2018-02-02 to 2018-02-26
Are conditions shown on logger consistent with ambient observations? | bar: 2018-02-02 to 2018-02-26
BAM-1020 PM10 Monitors
Arrival time: | 10:40 | 11:20 | 10:45 | 10:05 | 11:00 | 7:55 | 10:00
Any status error flags showing? | bar: 2018-02-02 to 2018-02-26
Are all monitor dates/times correct within 60 sec? | bar: 2018-02-02 to 2018-02-26
Are all filter transport systems operational? | bar: 2018-02-02 to 2018-02-26
Are all tape reel cover knobs secure? | bar: 2018-02-02 to 2018-02-26
Are hourly data being collected? | bar: 2018-02-02 to 2018-02-26
Is sampler inlet and downtube clean? | bar: 2018-02-02 to 2018-02-26
Is water collection bottle clean? | bar: 2018-02-02 to 2018-02-26
Are O-rings in good condition? | bar: 2018-02-02 to 2018-02-26
Comments section
2/20, Performed BAM QC checks.
2/23, Changed out BAM filter tapes. Cleaned rollers, capstans, etc.
1/29, Adjusted temperature on the ST-N BAM from -4.6 to -2.8 deg C.
Site ID List: | 2018-02-06: MAINTENANCE LIST FOR METEOROLOGICAL STATION
ST-NW = Northwest BAM site | 2018-02-06: Each Visit:; 2018-02-13: Inspect sensors and cables for signs of deterioration or damage and replace as needed
ST-N = North BAM site | 2018-02-06: Monthly:; 2018-02-13: Inspect temperature aspirator and remove dust accumulation
ST-NE = Northeast BAM site (Partisol samplers) | 2018-02-13: Inspect tower for plumb, and verify that all fasteners are intact and tight
ST-M = Meteorological Monitoring Site | 2018-02-06: Biannually:; 2018-02-13: Perform calibration check on sensors
ST-CE = Center East BAM site | 2018-02-13: Check wind sensor bearings
ST-SW = Southwest BAM site | 2018-02-06: Annually:; 2018-02-13: Replace wind sensor bearings
MAINTENANCE LIST FOR BAM-1020
|  | Each Visit: | Inspect filter tape and replace when needed (~2 months/roll; run SELF-TEST function in the TAPE menu).
Inspect filter tape for pinholes and irregular PM deposits
|  | Monthly: | Perform flow rate and leak check
Clean BAM nozzle and vane
Clean inlet hardware, capstan shaft and pinch roller
|  | Quarterly: | Verify BAM settings file
Flow system calibration
Disassemble and clean inlet
|  | Biannually: | Check pump muffler and tubing and replace as needed
Test filter RH and Temperature sensors
Test smart heater function
|  | Annually: | Perform zero background test
Clean internal debris filter
Remove and check membrane span foil
Beta detector count rate and dark count test
Rebuild vacuum pump
Replace nozzle o-ring