## MetBAM Checklist South
KCBX TERMINALS CO SOUTH |  | SITE VISIT DATE
| 2019-04-02 | 2019-04-05 | 2019-04-09 | 2019-04-12 | 2019-04-16 | 2019-04-18 | 2019-04-23 | 2019-04-26 | 2019-04-30
Meteorological Station
Arrival time: | 9:30 | 9:35 | 9:20 | 9:15 | 8:50 | 9:05 | 9:55 | 9:15 | 9:10
Is the temperature aspirator fan running? | bar: 2019-04-02 to 2019-04-30
Are all signal cables secured and intact? | bar: 2019-04-02 to 2019-04-30
Are wind vane/ anemometer cups intact and operational? | bar: 2019-04-02 to 2019-04-30
Is datalogger date/time correct within 60 sec? | bar: 2019-04-02 to 2019-04-30
Are conditions shown on logger consistent with ambient observations? | bar: 2019-04-02 to 2019-04-30
BAM-1020 PM10 Monitors
Arrival time: | 9:30 | 9:35 | 9:20 | 9:15 | 8:50 | 9:05 | 9:55 | 9:15 | 9:10
Any status error flags showing? | bar: 2019-04-02 to 2019-04-30
Are all monitor dates/times correct within 60 sec? | bar: 2019-04-02 to 2019-04-30
Are all filter transport systems operational? | bar: 2019-04-02 to 2019-04-30
Are all tape reel cover knobs secure? | bar: 2019-04-02 to 2019-04-30
Are hourly data being collected? | bar: 2019-04-02 to 2019-04-30
Is sampler inlet and downtube clean? | bar: 2019-04-02 to 2019-04-30
Is water collection bottle clean? | bar: 2019-04-02 to 2019-04-30
Are O-rings in good condition? | bar: 2019-04-02 to 2019-04-30
Comments section
4/2, Cleaned PM10 heads.
4/2, Changed time on the ST-SW BAM from 10:17:05 to 10:16 CST.
4/12, Performed BAM QC checks.
4/16, Inspected enclosure HVAC filters.
4/26, Could not perform BAM QC checks due to a DeltaCal problem.
4/26, Changed time on the ST-SE BAM from 09:51:05 to 09:50 CST.
4/30, Performed BAM QC checks.
Site ID List: | 2019-04-05: MAINTENANCE LIST FOR METEOROLOGICAL STATION
ST-NW = Northwest BAM site | 2019-04-05: Each Visit:; 2019-04-09: Inspect sensors and cables for signs of deterioration or damage and replace as needed
ST-N = North BAM site | 2019-04-05: Monthly:; 2019-04-09: Inspect temperature aspirator and remove dust accumulation
ST-NE = Northeast BAM site (Partisol samplers) | 2019-04-09: Inspect tower for plumb, and verify that all fasteners are intact and tight
ST-M = Meteorological Monitoring Site | 2019-04-05: Biannually:; 2019-04-09: Perform calibration check on sensors
ST-CE = Center East BAM site | 2019-04-09: Check wind sensor bearings
ST-SW = Southwest BAM site | 2019-04-05: Annually:; 2019-04-09: Replace wind sensor bearings
MAINTENANCE LIST FOR BAM-1020
|  | Each Visit: | Inspect filter tape and replace when needed (~2 months/roll; run SELF-TEST function in the TAPE menu).
Inspect filter tape for pinholes and irregular PM deposits
|  | Monthly: | Perform flow rate and leak check
Clean BAM nozzle and vane
Clean inlet hardware, capstan shaft and pinch roller
|  | Quarterly: | Verify BAM settings file
Flow system calibration
Disassemble and clean inlet
|  | Biannually: | Check pump muffler and tubing and replace as needed
Test filter RH and Temperature sensors
Test smart heater function
|  | Annually: | Perform zero background test
Clean internal debris filter
Remove and check membrane span foil
Beta detector count rate and dark count test
Rebuild vacuum pump
Replace nozzle o-ring