## General
KCBX SITE LOG
Date | Activity
2014-02-17 | 08:00 G. Mazik
Remove zero air scrubbers from BAM monitors that have
been running since 2/14 and replace with PM10 heads.
Perform site check. Everything OK.
2014-02-19 | 08:00 G. Mazik, P. Hight
Checked on NT and ST Partisol samplers. All running OK.
Removed shipping boxes from the NT and ST and placed in a rental
storage space.
2014-02-27 | 12:00 G. Mazik
North Terminal.
Perform routine site check,change background concentrations
on BAM monitors, and check on temperature difference status.
NT-NE was changed to -.0039 at 12:00 CST.
NT-SE was changed to -.0006 at 12:20 CST.
NT-NW was changed to -.0009 at 12:35 CST.
NT-SW was changed to -.0021 at 12:44 CST.
South Terminal.
ST-NE was changed to -.0033 at 13:00 CST.
ST-N was changed to -.0048 at 13:04 CST.
ST-NW was changed to -.0040 at 13:13 CST.
ST-SW was changed to -.0034 at 13:24 CST.
ST-SE was changed to -.0050 at 13:31 CST.
The ST met station was surrounded by ice.
All other checks were OK.
Several of the BAM monitors had the DT status changed.
ST Partisols did not run as scheduled on 2/25. Called Thermo for help.
Thermo is aware of the problem, but does not have a technical solution
at this time. They have received complaints from numerous other users
of this equipment during the recent cold temperatures.
2014-02-28 | 11:00 G. Mazik
J. Clarke reported that the South Terminal Partisols did not run as
scheduled on 2/28. Found the same problem as before, cassettes
were getting jammed in between the sample chamber and collection
tube. Called Thermo for help. Spoke with Sandy (508) 553-1559,
Sandy used to assemble these units and seemed more knowledgeable
than previous technical assistance staff I have spoken with.
Thermo instructed me to remove the shuttle cover
and adjust one of the cassette moving mechanisms. Did so on both
on both samplers. Ran 6 cassettes through on both samplers without
any problems. Set Partisols up to run on 3/1 as a makeup for 2/28.
Thermo reiterated that other users of the Partisols have had similar
problems during these extreme cold temperatures.
2014-03-02 | 07:00 G. Mazik
Dan from KCBX agreed to meet me this morning as an escort around the
South and North Terminals. The laboratory delivered the new batch of
metals and carbon cassettes late, so I unable to install them on 2/28.
South Terminal. Manually remove makeup samples that ran on 3/1.
Install cassettes to run on 3/3. North Terminal. Install next series of
cassettes into the Partisols. NT Partisols have been running without
incident.
2014-03-03 | 15:00 G. Mazik and P. Hight
South Terminal.
The 3/1 makeup ST Partisol samples ran OK. However, both
Partisols did not run on 3/3 due to a cassette jam in between
the sample chamber and collections tube. Reset both samplers to
run on 3/4 as a makeup.
2014-03-05 | 10:30 G. Mazik
South Terminal.
In to check on Partisols that ran on 3/4, as a makeup for 3/3.
Makeup samples ran on 3/4. Cassettes still jammed in between the
sample chamber and collection tube on both samplers, but barely.
We may be close to solving this problem.
Placed 3/4 cassettes into collection tubes. Removed 3/6
cassettes from sample chambers and placed in supply tubes to avoid
potential jams.
2014-03-07 | 11:00 G. Mazik
Perform site checks on the north and south terminals.
North Terminal.
All BAMs are working without incident. Several paper tapes will need
to be changed out next week. Partisols are working well.
South Terminal.
BAMs are working great. I could not access the ST-SW BAM due to
an active train that was blocking access. Partisols
still not working correctly. Removed filters that ran on 3/6 and got stuck.
Set the Partisols up to run on 3/9.
2014-03-11 | 10:30 G. Mazik
Perform QC checks on BAM and Partisol monitors.
There was no power at NT-SE. Notified M. Estadt of KCBX. In the
afternoon power was restored after about 6 hours being out.
Both Partisol monitors at the ST appear to working properly. Cassettes
did not get jammed advancing. Will hope that these samplers will
perform better with the warmer weather.
Removed Partisol cassettes that ran on 3/3, 3/6, and 3/9.
Site conditions are very muddy.
2014-03-12 | 13:00 G. Mazik
Emergency site visit. J. Clarke left a message that all four wind sensors
appear frozen. There had been a snow event with high winds last night.
Called M. Estadt to arrange for a bucket truck to take corrective action.
Found WS and WD sensors covered with ice and snow.
Work was performed by the bucket truck operator under my direction.
With a combination of deicing spray and light tapping, removed ice
and snow from all four sensors, NT wind speed cups were turning,
but were filled with snow and ice, and not collecting valid data.
Valid data collection resumed at 14:00 CST.
The initial estimate is that problems began at approximately 04:00.
The top of most surface areas are frozen.
2014-03-14 | 12:00 G. Mazik
Replace filter tape on BAMs NT-NW, NT-SW, and NT-SE. All were
running low. Installed new shipment of Partisol filters from RTI.
Ground areas are very wet and the mud is deep. The south facing piles
at the NT seem to be drying out. [Note: Partisol samplers at the NT ran only 22 hours on
March 12 due to a power failure, most likely related to an ice storm. This was noted on the
lab CoC form. J. Clarke 4/25/14.]
2014-03-17 | 11:00 G. Mazik
Perform QC checks on BAM and Partisol monitors. Clean BAM and
Partisol PM10 sampling heads.
Removed Partisol cassettes that ran on 3/12 and 3/15. The ST metals
sampler jammed while loading the 3/15 cassette, so there is no sample
for that date. Conditions are still very wet.
2014-03-19 | 08:45 G. Mazik, J. Sorkness
J. Sorkness is on site to perform audits of the particulate samplers and
meteorological systems. There is a consistent rain which is causing
delays with the audits. There is at least four inches of standing water
surrounding the ST meteorological monitoring station. Completed the
ST met audit. Went to the NT to audit that station. Performed audits
on the NT BAMs and Partisols. All OK.
11:00 D. Szatkowski
The BAM units at the NW and SE site were put into the setup mode around 11AM and were mistakenly not taken out of the setup mode until 03/20/14 around 11AM. Lost about a days worth of data. The BAM's at the south terminal were auditedon 03/20/14. All the BAM's at both sites passed the audit. All the BAM's were reporting concentrations in actual mode.
This parameter was changed on 03/20/14 such that the systems will report concentrations in standard mode. All the concentration data
collected up to 03/20/14 at 13:00, will have to be corrected to report standard conditions.
2014-03-20
08:30 G. Mazik, J. Sorkness
Perform audits of the BAMs and Partisols at the ST. All OK. J. Clarke
notified us that the NW and SE BAMs at the NT were not placed in
sample mode following yesterdays audits. Returned to the NT to place
those two BAMs in sample mode. Also, changed the concentration
mode on the BAMs from actual to standard. Change times are as
follows:
ST-NE 11:25 CST
ST-SE 11:30 CST
ST-SW 12:00 CST
ST-N 12:10 CST
ST-NW 12:15 CST
NT-NW 12:30 CST
NT-SW 12:25 CST
NT-NE 12:38 CST
NT-SE 12:43 CST
2014-03-26
11:00 G. Mazik
Performed site check at the NT. All OK. BAM NT-NE tape will need to
be changed during my next site visit. For some reason, both Partisols
advanced a filter before running on 3/21. I will put both 3/21 filters back
into the supply tubes to run on 3/30.
At the ST, all instruments were operating normally, except for the metals
Partisol. Filter cassettes were jammed again. We lost the 3/21 and
3/24 filters on that sampler. Will adjust both ST Partisols again using
directions from Thermo that was supposed to solve this recurring
problem. After adjustment, installed ST Partisol filters to run on 3/27 and
3/30. Found the same problem as the NT, 3/21 were advanced without
running. Not sure why this happened on all four Partisols following the
audit. I will contact Thermo to find out if there was
something we did during the audit that may have caused this.
Ground was wet. Presence of fugitive dusk not observed.
2014-03-28
11:15 G. Mazik
NT - Installed next batch of cassettes into the Partisols (through 4/14).
Replaced the NT-NE BAM filter tape.
ST - Installed next batch of cassettes into the Partisols. Both samplers
ran without incident on 3/27. Will hope the 3/26 adjustments did the
trick, once and for all. Observed occasional fugitive dust blowing
near the ST-NE site.
2014-03-31
8:45 D. Szatkowski
Tape error is being reported at the NT-NW site. It went off line Friday, 03/28/14. The winds
were coming out of the South at 10mph.Our back-up field operator, Phil Hight, was on site
about noon on Monday, 3/31/14 and simply had to tighten the right side spool nut. The
system was backon-line shortly after the noon hour. The operator noted that the site
conditions were quite wet and muddy at the time of the repair and that there were no visible
signs dust plumes. He also mentioned that the site had received recent rainfall.
2014-04-01
08:30 D. Szatkowski
On site to install new data communications hardware that will improve
data collection from all monitors.
2014-04-02
08:30 D. Szatkowski
Continue installing data communications hardware.
Communications hardware installation complete. Notified G. Mazik that
filter tapes on several of the BAMs at the ST will need replacing next
week.
2014-04-04
10:30 G. Mazik
Perform routine site check. All instruments operating as designed.
BAM NT-SE was very difficult to get to. Water level is at least a foot
in most areas. Got stuck a couple of time in the 18" deep mud, but was
able to work my way out. For some reason, water is pooling in the area.
Checked on the operation of the ST-metals Partisol. Seems to be
working well after the last adjustment. Ground was very wet after recent
rains. The KCBX road cleaning crews are finally getting down to
pavement, after weeks of scraping layers of frozen product.
2014-04-08
10:30 G. Mazik
In to perform QC checks on the BAMs. Found all to be operating within
project limits. Was unable to check BAM NT-SE. It is still surrounded
deep water and mud. Will try again on 4/10 to see if the water level has
gone down. If not, I will try to approach from along the south property
line. Piles and roads are really starting to dry out. There was a huge
cargo vessel being loaded near BAM ST-SW. Base on size, it will likely
take several days to load. Paved roads I observed were clean and wet.
Replaced paper tapes on all ST BAMs. Change times are listed below:
ST-NW 10:05
ST-N 10:25
ST-NW 10:55
ST-SE 11:15
ST-NE 11:35
2014-04-10 | 12:00 G. Mazik
Perform QC checks on the Partisols. All were found to be operating within project limits
except for the leak check on the ST-metals monitor. Initial reading was 26 vs. the allowable
25. Cleaned the seals and retested. Second reading was 5. Still having problems with
access to the NT-SE BAM. Water surrounding it seems higher today. Will contact
J. King about the situation to see if water levels can be lowered.
Water trucks keeping the roads wet. ST street sweepers keeping the paved roads clean.
Did not observe fugitive dust at either terminal.
Collected Partisol cassettes that ran on 4/2, 4/5, and 4/8.
Peformed site checks at both terminals. All OK.
2014-04-15 | 13:00 G. Mazik
Perform site checks at the NT. All OK. Performed a QC check on the NT-SE
BAM. This site is still very difficult to get to. The return walk from the site was easier
because I was able to walk on a paved area next to the river. Usually, this area is not
accessible. However, work crews were installing sand bags along the edge of the river
and the fence gates were unlocked. NT-SE BAM was offline from 13:40-13:50.
Collected Partisol cassettes that ran on 4/11 and 4/14.
Installed the next batch of Partisol cassettes that will run through 4/29.
2014-04-18 | 10:00 G. Mazik
Found KCBX closed for business for the the Good Friday holiday. Should have checked
with them first . Was planning on replacing the air conditioner at the NT-NW site and
performing site checks at the ST.
2014-04-22 | 10:00 G. Mazik
Installed a replacement air conditioner at the ST-NW. Called Ben to verify that it was
working correctly. Perform site checks at the ST. Perform QC checks on the ST BAMs.
The BAMs and other monitoring equipment was found to be operating correctly.
Found the 4/20 cassette jammed again on the ST-metals Partisol. Adjusted the shuttle
again using instructions from Thermo that is supposed reduce cassette jams.
Both ST Partisols are set up for the 4/23 run date.
2014-04-25 | 09:00 G. Mazik
Perform QC checks on the ST Partisols. Perform QC checks on the NT Partisols and NT
BAMs. All samplers were found to be operating properly. Conduct a site check at the NT.
Collected Partisol cassettes that ran on 4/17, 4/20, and 4/23.
2014-04-30 | 10:15 G. Mazik
Collect Partisol filters that ran on 4/26 and 4/29.
Install Partisol filters for the period 5/2 through 5/14.
Steady rain falling. Will delay routine cleaning of BAM nozzles and vanes until 5/2.
2014-05-02 | 07:40 G. Mazik
On site to clean nozzles and vanes of the BAM monitors. Results are listed below:
ST-N Time=07:40 CST As found 0.5, as left 0.4.
ST-NW Time=08:30 CST As found 0.5, as left 0.3.
ST-NE Time=09:00 CST As found 0.6, as left 0.5.
ST-SW Time=10:00 CST As found 0.6, as left0.4.
ST-SE Time=10:30 CST As found 0.5, as left 0.4.
NT-SW Time=11:30 CST As found 0.6, as left 0.4.
NT-NW Time=11:45 CST As found 0.6, as left 0.4.
NT-NE Time=12:30 CST As found 0.5, as left 0.4.
NT-SE Time=12:50 CST As found 0.7, as left 0.5.
2014-05-06 | 10:15 G. Mazik
Perform site checks at the ST. All OK.
Perform QC checks on the ST BAMs, except for ST-SW. There is a train blocking the way.
Conduct site checks at the NT. All equipment is operating properly. Found the 5/5 cassette
jammed above the collection tube of the NT-Metals Partisol. The 5/8 cassette was properly
located in the sample chamber. I removed it, and placed it back into the supply tube just to
be safe. The sample dates and cassette serial numbers were listed correctly in the menu.
This is the first time I have experienced this issue at the NT.
2014-05-09 | 12:00 G. Mazik
Perform QC checks on the ST-SW BAM. Perform ST site checks. All OK.
Conduct QC checks on the ST and NT Partisol samplers. Perform QC checks on the
NT BAMs. All BAMs and Partisols were found to be operating within allowable project
limits. Conduct NT site checks. All OK.
2014-05-13 | 09:10 G. Mazik
Perform site checks. OK.
Conduct cleaning of PM10 sampler heads. Cleaning times are listed below.
NT-NE 09:10 CST
NT-SE 09:30 CST
NT-SW 09:55 CST
NT-NW 10:05 CST
ST-NW 10:30 CST
ST-N 10:40 CST
ST-SW 10:55 CST
ST-SE 11:00 CST
ST-NE 11:10 CST
Changed ST-NW BP multiplier from 290 to 300.
Collected Partisol cassettes that ran on 5/2, 5/5, and 5/8.
2014-05-16 | 12:00 G. Mazik
In to investigate several intermittent pressure readings for the ST-NE BAM. J. Clarke had
reported this issue in the afternoon of 5/15. Conducted the following while checking out
the ST-NE BAM. Checked the control cable to the pump, trimmed and reseated the flow
tube to the pump, tightened sample flow tube seal, and tightened the enclosure door
latches. Conducted a QC check on the unit and found everything operating within
project limits. Reported results to D. Szatkowski. Since the BAM had passed the QC
checks and there have been no pump issues since 5/15, we leave the existing BAM
in place and intend to keep a close eye out for recurrence.
Performed site checks on the ST and NT instrumentation. All OK.
Removed Partisol cassettes that ran on 5/11 and 5/14.
Changed NT-NW BAM filter tape at 15:15 CST.
2014-05-20 | 09:20 G. Mazik
Perform QC checks on the BAM monitors. All measurements were within allowable project
limits except for ambient temperature on the NT-SE unit, which was reading high. Adjusted
the temperature on that BAM. Conducted site checks of the remaining Partisol and
meteorological equipment. OK.
2014-05-22 | 07:50 G. Mazik
Perform QC checks on the Partisol samplers. The four Partisol units were found to be
operating with allowable limits.
Changed NT-NE BAM filter tape at 07:53 CST.
Changed NT-SE BAM filter tape at 09:10 CST.
Changed NT-SW BAM filter tape at 09:30 CST.
Removed Partisol cassettes that ran on 5/17 and 5/20.
Completed site checks at the NT and ST. All normal except for the ST-SW BAM, which
was found with no power at approximately 11:00 CST. I immediately notified M. Estadt and
D. Szatkowski of the situation. An electrician from Shaleen Electric found that power to
the BAM enclosure was turned off during a Lockout/Tagout operation for pumps that are
on the same circuit. Power was restored and the ST-SW BAM became operational at
12:15 CST.
2014-05-28 | 08:05 G. Mazik
J. Clarke called to report that data communication to the ST meteorological system was
not working. D. Szatkowski suggested resetting the router and reinstalling the
air card, in case of a bad connection. The system came back online and URS was able
to retrieve the data.
Performed site checks at the ST and NT. All OK.
Cleaned the NT BAM nozzles and vanes. Leak check results are listed below.
NT-NE Time=10:55 CST As found 0.5, as left 0.4.
NT-SE Time=11:14 CST As found 0.8, as left 0.5.
NT-SW Time=11:41 CST As found 0.6, as left 0.4.
NT-NW Time=11:55 CST As found 0.5, as left 0.5.
Rechecked the NT-SE ambient temperature. BAM=11.5. Ref=12.8.
2014-05-30 | 08:55 G. Mazik, P. Hight
Performed ST and NT site checks. All running as expected.
Clean nozzles and vanes of the ST BAM monitors. Results are listed below:
ST-SE Time=09:14 CST As found 0.5, as left 0.5.
ST-N Time=09:34 CST As found 0.5, as left 0.5.
ST-NE Time=09:55 CST As found 0.7, as left 0.5.
ST-SW Time=10:25 CST As found 0.5, as left 0.5.
ST-NW Time=12:03 CST As found 0.4, as left 0.3.
Filter tapes on all ST BAM monitors were changed after cleaning and leak checks.
2014-06-04 | 09:00 P. Hight
In to investigate a communications problem with the ST meteorological station. Rebooted
the router. Communications were restored. No data lost.
2014-06-06 | 09:20 G. Mazik
Perform site checks at the ST and NT. All OK except for two BAM times.
The time on the ST-NW BAM was changed from 09:36 to 09:34 CST.
The time on the ST-NE BAM was changed from 09:53 to 09:51 CST.
The NT-SE BAM display did not show "Normal". Changed this BAM to "Normal" mode.
The NT-SE BAM seemed to be operating correctly.
Removed Partisol cassettes that ran on 6/1 and 6/4.
2014-06-09 | 07:35 J. Sorkness, G. Mazik
J. Sorkness onsite to perform audits of the BAM samples. BAM samplers were taken
offline at the following times:
ST-NW taken offline at 07:50 CST.
ST-N taken offline at 08:08 CST.
ST-SW taken offline at 08:25 CST.
ST-SE taken offline at 08:40 CST.
ST-NE taken offline at 08:55 CST.
NT-NE taken offline at 11:45 CST.
NT-SE taken offline at 12:45 CST.
NT-SW taken offline at 13:02 CST.
NT-NW taken offline at 13:15 CST.
2014-06-10 | 07:30 J. Sorkness, G. Mazik
J. Sorkness onsite to assist with quarterly calibrations of the meteorological systems.
Bucket truck and operator delayed until 09:30 CST due to a personnel change.
ST meteorological instruments were found to be operating within allowable limits.
NT meteorological instruments were found to be operating withing allowable limits.
2014-06-12 | 10:30 G. Mazik
Perform site check at the ST and NT. All equipment working properly.
Remove Partisol cassettes that ran on 6/7 and 6/10.
Installed a new batch of Partisol cassettes that will last until 6/28.
2014-06-16 | 08:00 G. Mazik
Replaced air conditioning filters with brand new ones on all enclosures except for ST-NW
and NT-NW. These two air conditioners use a different size filter. However, the ST-NW
and NT-NW filters were cleaned by hand, and reinstalled back into the air conditioning unit.
The time was changed on the NT-NE BAM from 09:57 to 09:55 CST.
The time was changed on the ST-SW BAM from 10:51 to 10:52 CST.
All Partisols were running as scheduled.
Removed Partisol cassettes that ran on 6/13.
Performed site checks at the ST and NT. All instruments were found to be operating
properly.
2014-06-19 | 08:30 G. Mazik
Retested the NT-NW ambient temperature sensor with BGI DeltaCal s/n 1377. Results
are listed below: Temperature checks were made at 09:00 CST.
Sensor Output DeltaCal
19.7 19.7
19.6 19.7
19.4 19.7
Performed QC checks on the NT and ST BAM monitors. All BAMs were operating within
allowable project limits. BAMs were taken offline at the following times:
NT-NW taken offline at 09:05 CST.
NT-SW taken offline at 09:20 CST.
NT-NE taken offline at 09:40 CST.
NT-SE taken offline at 09:50 CST.
ST-NW taken offline at 10:30 CST.
ST-N taken offline at 10:40 CST.
ST-NE taken offline at 10:45 CST.
ST-SE taken offline at 10:55 CST.
ST-SW taken offline at 11:35 CST.
Conducted site checks at the NT and ST. Instrumentation was operating properly.
All Partisols were running as scheduled.
2014-06-20 | 10:30 G. Mazik
Onsite to investigate data communications problems with the ST meteorological station.
At the request of D. Szatkowski, I switched the LN-100 network interface at the ST with the
one at the NT. Waited for approximately two hours to see if the replacement LN-100
would improve communications at the ST meteorological station. D. Szatkowski indicated
that communications with the ST meteorological station were good.
2014-06-24 | 08:00 G. Mazik
Perform QC checks on the NT and ST Partisol samplers. All samplers were found to be
operating within allowable limits.
Met with the electrical contractor at 12:00 CST to change out the ST meteorological
system. Change out was completed at 13:30 CST.
Conduct site checks at the NT and ST. All OK.
2014-06-27 | 07:20 G. Mazik
Perform site checks at the NT and ST. All instrumentation was operating correctly.
AT 9:00 CST, met P. Hight at Public Storage to move equipment boxes from the current
storage space to a larger one that will accommodate spare equipment.
2014-07-01 | 09:00 G. Mazik
Perform site checks at both terminals. All OK.
Removed Partisol cassettes that ran on 6/25 and 6/28.
Investigated low pressure warning at the ST-NE BAM.
Did not observe water inside the enclosure or other potential issues. It had
rained hard with high winds earlier in the morning. Performed flow and leak checks.
Flow = 16.62. Leak=0.5.
Found 1/2 inch of water in the ST-NW BAM jar.
2014-07-08 | 09:00 G. Mazik, P. Hight
Moved spare Partisol and air conditioner unit from KCBX property to URS's offsite storage
facility.
Perform QC checks on the ST BAMs. Clean ST BAM PM10 heads and nozzles. Repeat
leak checks on the ST BAMs. As found and as left leak check results are listed below.
ST-SW taken offline at 10:05 CST. Leak check as found 0.3, as left 0.4.
ST-SE taken offline at 10:30 CST. Leak check as found 0.3, as left 0.4.
ST-NE taken offline at 12:30 CST. Leak check as found 0.6, as left 0.5.
ST-NW taken offline at 13:00 CST. Leak check as found 0.3, as left 0.3.
ST-N taken offline at 13:20 CST. Leak check as found 0.4, as left 0.4.
Conduct site checks at the ST and NT. All OK.
Collect Partisol cassettes that ran on 7/1, 7/4, and 7/7.
2014-07-09 | 10:15 G. Mazik, P. Hight
In to investigate an intermittent cellular signal strength issue at the ST-SE BAM.
Replaced the antenna, main connection cable, and the adapter cable. D. Szatkowski
indicated that the signal strength has improved following these changes.
The ST-SE BAM data are invalid starting at 10:15 CST until 11:00 CST due to the
ambient temperature probe being disconnected during the change out.
2014-07-11 | 10:45 G. Mazik
Conduct site checks at the ST. All OK.
Perform QC checks on the NT BAMs. Clean NT BAM heads and nozzles. Repeat
leak checks on the NT BAMs. As found and as left leak check results are listed below.
NT-NE taken offline at 11:40 CST. Leak check as found 0.3, as left 0.4.
NT-SE taken offline at 12:20 CST. Leak check as found 0.8, as left 0.5.
NT-SW taken offline at 13:10 CST. Leak check as found 0.6, as left 0.4.
NT-NW taken offline at 13:30 CST. Leak check as found 0.6, as left 0.5.
Ambient temperature on NT-SE BAM was adjusted from 22.0 to 25.2.
Conduct site checks at the NT. All OK.
Observed an welding generator near the NT-SW BAM. D. Szatkowski advised KCBX
of the situation.
2014-07-15 | 08:40 G. Mazik
Perform site checks at the ST and NT. All OK except for time changes on the ST-SW and
NT-SW BAMs. The ST-SW BAM was taken offline at 09:30 CST and the NT-SW BAM
was taken offline at 11:35 CST.
Removed Partisol cassettes that ran on 7/10 and 7/13.
2014-07-18 | 08:30 G. Mazik
Conduct QC checks of the ST and NT Partisol samplers. The four samplers were found
to be operating within allowable project limits. The NT-carbon Partisol time was changed
from 11:51 to 11:52 CST.
Perform site checks at the ST and NT. All OK.
2014-07-22 | 09:00 G. Mazik
Perform site checks of the ST and NT monitoring equipment. All OK except for both
NT Partisols. The 7/19 NT-carbon cassette was jammed. Reset the 7/22 sample to run
on 7/23. The 7/22 cassette was jammed on the NT-metals Partisol. Reset that sample
to run on 7/23.
Removed Partisol cassettes that ran on 7/16 and 7/19.
2014-07-24 | 09:00 G. Mazik
Perform QC checks on the ST and NT BAM monitors. All parameters were found to be
operating within acceptable ranges except for ambient temperature at the NT-SE
monitor. This temperature was reading about 3 degrees low. Adjusted the NT-SE
temperature from 17.2 to 20.0 using the DeltaCal s/n 1377.
Replaced or cleaned all enclosure air conditioning filters.
2014-07-28 | 09:00 G. Mazik
Perform ST and NT site checks. All OK.
Replace NT wind speed sensor. Old sensor s/n was P15942 and the new one is F6198.
The wind speed sensor was replaced at 13:15 CST.
All Partisols are running as scheduled.
2014-07-29 | 10:30 G. Mazik
Power was lost at the ST-SW BAM from 08:00 to 11:00 CST due to a Lockout/Tagout
operation.
2014-08-01 | 08:00 G. Mazik
Remove Partisol cassettes that ran on 7/22, 7/23(NT), 7/25, 7/28, and 7/31.
Found filter ST-073114-CR jammed in between the supply tube and sample chamber.
Install the next batch of Partisol cassettes that will run though 8/15.
Replace BAM tapes at the ST. The BAMs were taken offline at the times listed below.
ST-NW offline at 11:00 CST.
ST-N offline at 11:15 CST.
ST-SE offline at 11:25 CST.
ST-NE offline at 11:35 CST.
ST-SW offline at 12:45 CST.
2014-08-05 | 09:30 G. Mazik
Performed site checks at the ST and NT. All OK except for the time on the NT-NW and
ST-NW BAMs. The time on the NT-NW BAM was changed from 10:34 to 10:34 CST and
the ST-NW BAM from 11:26 to 11:25 CST.
Found the ST-N running, but not in sample mode at 11:20 CST. Placed it back in sample
mode. Upgraded the ST-N data logger at 13:30 CST.
2014-08-06 | 09:30 G. Mazik, D. Szatkowski
The remaining four data loggers at the ST were upgraded 8/6/14 between the hours of 10:00
and 11:00 CST. The remaining data loggers at the NT were upgraded 8/6/14 between the
hours of 11:00 and 12:00 CST. The NT-NW data logger was upgraded at 14:00 CST.
When we verified the operation of the ambient temperature at the NT-NW site, we
discovered that the temperature was reading about 3 deg C lower than the Delta Cal std.
The entire 12VDC aspirator and temperature sensor was swapped out with the onsite spare
unit. This corrected the problem. Most likely the temp sensor has drifted out of spec.
Looking back at the data base, this temperature drift began to occur on 6/26/14. We may
need to invalidate the ambient temperature for this time block and replace it with the
temperature from the ST-MET station. There should have been minimum interruption in the
data in the RTMC application. We missed approximately an hour of data at various times
in the web site data base while the loader scripts were modified by the database manager.
Overall the upgrade went quite smooth. We will monitor the data collection closely for the
next couple to verify smooth operation.
The ST-SW data logger was replaced at 08:35 CST.
The ST-SE data logger was replaced at 08:50 CST.
The ST-NE data logger was replaced at 09:00 CST.
The ST-NW data logger was replaced at 09:20 CST.
The NT-NE data logger was replaced at 10:15 CST.
The NT-SE data logger was replaced at 10:30 CST.
The NT-SW data logger was replaced at 10:50 CST.
The NT-NW data logger was replaced at 12:15 CST.
2014-08-08 | 09:30 G. Mazik
Perform QC checks on the NT and ST BAMs. All units were found to be operating within
allowable limits.
Changed out NT-NW power and signal wires with longer ones which make it easier to
place them into the enclosure. As a result, the ambient temperature recorded at the
NT-NW met station are invalid for the hours 09:00-10:00 CST. The temperature before
the wiring change was 24.9 deg C, and after the change was 25.3 deg C. The temperature
was check with the Delta Cal.
Conducted site checks at the NT and ST. All OK.
Removed Partisol cassettes that ran on 8/3 and 8/6.
2014-08-11 | 09:30 G. Mazik
Perform site checks at the NT and ST. All OK.
2014-08-14 | 10:30 G. Mazik
Conducted QC checks on the NT and ST Partisol samplers. All parameters were found
to be operating within project limits except for sampler times. Sampler times were running
from 1 to 2 minutes slow. Times were adjusted using the reference time.
Collected Partisol cassettes that ran on 8/9 and 8/12.
2014-08-18 | 08:20 G. Mazik
An alarm condition at the ST-SE BAM was reported at 07:30 CST. This alarm was caused
a momentary power outage. The BAM was placed back in Operate/Normal mode at
08:30 CST.
Installed a UPS at the NT-NW enclosure at 10:45 CST. Site was back online at 11:00 CST.
Adjusted the ambient temperature on the NT-SW BAM from 20.9 to 23.7 degrees C.
The NT-SW BAM was taken offline at 11:10 CST and returned to service at 11:30 CST.
Conducted site checks at the NT and ST. All OK.
Removed the Partisol cassettes that ran on 8/15 and 8/15 Blank.
2014-08-22 | 07:40 G. Mazik
Performed QC checks on the NT and ST BAMs, except for ST-N. All tested BAMs were
found to be operating within allowable project limits except for time on the NT-NW BAM.
Time was adjusted from 08:44 to 08:46 CST on the NT-NW BAM.
Observed that there was no power at the ST-N BAM at 09:05 CST. Notified D. Szatkowski,
J. Clarke, and M. Estadt of the problem. M. Estadt advised that heavy rains overnight
caused the Diesel Shop to flood and the circuit breakers to trip. The ST-N BAM is on the
same circuit as the Diesel Shop. KCBX began pumping out the Diesel Shop and the area
surrounding it with pumps and vacuum trucks. Water pumping and collection was
completed at 16:45 CST. The power was initially restored at 17:05 CST, but the circuit
breaker to the ST-N tripped a couple of times. The electrician made adjustments and
power was restored at 17:15 CST. I observed the BAM operation until 17:30 CST after
placing it into Operate/Normal mode. There were no error messages and the unit
seemed to be back to normal. The power outage appears to have begun at 08:00 CST.
All other instrumentation was operating properly during the NT and ST site checks.
Removed Partisol cassettes that ran on 8/18 and 8/21.
2014-08-26 | 09:00 G. Mazik
D. Szatkowski reported that the NT met temperature response was fluctuating a degree or
two. Inspected the signal and aspirator wire connections. Did not find a loose wire. But
removed and reattached the wires in the aspirator and on the data logger. Checked the
temperature output against the Delta Cal standard. Data logger T=27.2 the Delta Cal=27.6.
Performed a leak check on the NT-NW BAM at 09:30. Leak as found=0.6. As left=0.5.
Performed a leak check on the NT-SW BAM at 10:00. Leak as found=0.6. As left=0.4.
Performed a leak check on the NT-NE BAM at 11:30. Leak as found=0.3.
Performed a leak check on the NT-SE BAM at 11:50. Leak as found=0.7. As left=0.5.
Performed a leak check on the ST-NW BAM at 12:30. Leak as found=0.3.
Conducted a QC check on the ST-N BAM at 12:45. OK.
Performed a leak check on the ST-SW BAM at 13:05. Leak as found=0.4.
Performed a leak check on the ST-SE BAM at 13:15. Leak as found=0.4.
Did a temperature check the ST met site at 13:25. Data logger=26.7. Delta Cal=26.3.
Performed a leak check on the ST-NE BAM at 13:30. Leak as found=0.7. As left=0.4.
2014-08-28 | 08:00 J. Sorkness, G. Mazik
J. Sorkness onsite to perform quarterly audits of the air quality and meteorological
instruments.
BAM monitors were taken offline at the times listed below for the audits.
ST-N BAM offline at 08:00 CST.
ST-NW BAM offline at 08:20 CST.
ST-SW BAM offline at 08:35 CST.
ST-SE BAM offline at 08:45 CST.
ST-NE BAM offline at 08:55 CST.
ST Partisols were audited beginning at 09:00 CST.
NT-NE BAM offline at 10:00 CST.
NT-SE BAM offline at 10:50 CST.
NT-SW BAM offline at 11:05 CST.
NT-NW BAM offline at 12:00 CST.
NT Partisols were audited beginning at 11:30 CST.
NT met station audits began at 13:00 CST. Wind speed and temperature were OK.
Wind direction failed the linearity check. The man lift operator had to leave to respond to
an emergency. Will retest the wind direction sensor tomorrow.
2014-08-29 | 07:00 J. Sorkness, G. Mazik
J. Sorkness onsite to perform quarterly audits of the meteorological instruments.
The ST met station was audited from 10:00 to 11:00 CST. Wind direction, wind speed,
and temperature were found to be operating within allowable limits.
Returned to the NT met station to retest the wind direction sensor from 11:10 to 11:45 CST.
The old sensor failed the linearity test again and was replaced. The new sensor passed
the audit.
2014-09-02 | 09:45 G. Mazik
Perform site checks at the NT and ST. All OK except for the time on the ST-N BAM. The
time on the ST-N BAM was changed from 11:25:30 to 11:27 CST. The ST-NW BAM was not
in Operate/Normal mode, but seemed to be functioning properly. Placed this BAM in
Operate/Normal mode at 11:35 CST. Cleaned the ambient temperature probe at the ST-met
station starting at 12:00 CST. Checked temperature following cleaning using the Delta Cal.
Delta Cal = 28.2, data logger = 27.9 deg C.
2014-09-04 | 08:00 G. Mazik
Onsite to investigate a communications problem at the NT-SW BAM reported by
D. Szatkowski. Checked for loose wires. None found. Reset the data logger.
Communications with the NT-SW BAM have been restored.
Perform site checks at the NT and ST. All OK.
2014-09-06 | 09:00 G. Mazik
At the ST to check on a loss of communications with the ST-met station. Found all
devices without power. Replaced the power supply and the lights on the network interface
and router came back on. However, the data logger and fan did not. Checked voltage to
the data logger and fan. Found 12.7 volts DC to each device. When inserting the
power plug in the data logger, the power light came on for a second then went blank.
Typically, the power light will flash every 15 seconds when operating properly.
Also, the display screen will show system information. The screen was blank. Spoke
with D. Szatkowski about the issue. D. Szatkowski requested that power be disconnected
to the station. A replacement data logger and fan will be sent out on Monday, 9/8.
There appears have been an electrical power surge to the ST-met station that caused
the damage to the instrumentation. It is unlikely that the power supply, data logger, and
fan would all fail at the same time without an external influence.
2014-09-09 | 09:00 G. Mazik
Investigate communications issue with the NT-SW BAM. Reset the data logger and
communications were restored to that station.
Perform QC checks on the NT and ST BAMs. All monitors were found to be operating
within allowable project limits. Changed time on the NT-NW BAM from 9:14:55 to 9:16 CST.
BAM monitors were taken offline at the times listed below for the checks.
ST-N BAM offline at 13:40 CST.
ST-NW BAM offline at 13:50 CST.
ST-SW BAM offline at 14:10 CST.
ST-SE BAM offline at 14:25 CST.
ST-NE BAM offline at 13:55 CST.
NT-NE BAM offline at 9:40 CST.
NT-SE BAM offline at 9:50 CST.
NT-SW BAM offline at 8:30 CST.
NT-NW BAM offline at 9:10 CST.
Begin work on the ST-met station at 10:30 CST.
Replaced the data logger at the ST-met station. The AT and BP channels starting
recording, but the WS and WD did not. The WS and WD sensors were probably
damaged during the power surge on 9/5, along with the power supply, data logger,
and enclosure fan. KCBX supplied a man lift for the sensor change out. Replaced the WD
and WS sensors at 12:45 CST. The ST-met station became operational at 13:00 CST.
Checked WD alignment with marks on the sensor. Obtained 178 degrees without a
hose clamp to hold the vane in place. Checked the temperature with the Delta Cal.
Delta Cal = 25.8 deg C, data logger = 26.0. Delta Cal BP = 746, data logger = 745.
Conducted site checks at the NT and ST. All OK.
Collected Partisol cassettes that ran on 9/2, 9/5, and 9/8.
2014-09-10 | 12:00 P. Hight
Onsite to investigate a problem with the ambient temperature on the NT-SE BAM. It has
been reading higher than ambient conditions. Checked the temp/BP probe plug.
It may have come loose during the recent high winds. Found it was properly attached.
2014-09-11 | 10:00 G. Mazik
Replaced the temp/BP sensor on the NT-SE BAM. The as found reading was
34.0 deg C, using the Delta Cal. As left temperature with the new sensor was 11.5 deg C.
Perform site checks at the NT and ST. All OK.
2014-09-16 | 09:00 G. Mazik
Perform QC checks on the ST and NT Partisol samplers. All OK except for the times that
were about one minute slow. Partisol times were corrected to the reference times.
Took the NT-SE BAM offline at 11:45 CST to turn on the delta-t function and to change the
time from 11:50:01 to 11:51 CST. Checked the AT of the NT-SE BAM. Delta Cal = 16.7,
BAM = 17.4 deg C.
Conducted site checks at the ST and NT. All equipment running normally.
Collected Partisol cassettes that ran on 9/11 and 9/14.
Arranged for a man lift to check the alignment of the NT-met WD sensor. Found the
crossarm alignment at 5 deg - 3.6 deg west declination = 1.4 deg (almost true north).
Checked the alignment with a survey point of 266 - 3.6 = 262.4 deg. Data logger = 268 deg.
I will do additional alignment checks, when the man lift operator is available.
2014-09-19 | 08:35 G. Mazik D. Szatkowski
Perform site checks at the NT and ST. All OK.
The man lift operator was not available today to check the WD alignment.
Data communications with the NT-SW BAM was lost at approximately 13:00 CST.
D. Szatkowski notified T. Chandler of KCBX, who referred him to Earl.
Earl visited the NT-SW site and found the power on. He later reset the data logger,
which restored communications to the site, and it resumed normal data collection.
2014-09-22 | 09:35 G. Mazik
Perform QC checks on the NT and ST BAMs. All monitors were found to be operating
within allowable project limits. Changed time on the NT-NE BAM from 10:04:05 to 10:03 CST.
BAM monitors were taken offline at the times listed below for the checks.
NT-NE BAM offline at 10:00 CST.
NT-SE BAM offline at 10:30 CST.
NT-SW BAM offline at 11:00 CST.
NT-NW BAM offline at 11:20 CST.
ST-NW BAM offline at 12:00 CST.
ST-N BAM offline at 12:20 CST.
ST-SW BAM offline at 12:40 CST.
ST-SE BAM offline at 13:05 CST.
ST-NE BAM offline at 13:30 CST.
Asphalt road paving was occurring to the south of the ST-N BAM.
Weed cutting was occurring near the NT-SW BAM.
Conducted a NT-met WD observation. Observed 90 deg., data logger 80-90 deg.
Conducted a ST-met WD observation. Observed 45 deg., data logger 30-50 deg. Winds
at the ST-met station were light and the vane was fluctuating.
Performed site checks at the NT and ST. All OK.
2014-09-25 | 16:30 G. Mazik
J. King reported that power would be turned off at the NT-NW BAM site for maintenance
by the local utility company (Commonwealth Edison) at 16:00 CST. The NT-NW began
collecting data at again at 17:00 CST. Hour 16:00 CST is invalid due to the power failure.
2014-09-26 | 08:35 G. Mazik
Performed site checks at the NT and ST. All OK.
Conducted QC checks on the NT-NW BAM following the power failure on 9-25-14. OK.
Checked the NT-NW BAM pump and pump hoses. OK.
Removed Partisol cassettes that ran on 9-23-14.
Installed the next batch of Partisol cassettes that will run through 10-14-14.
2014-10-03 | 10:30 G. Mazik
Performed site checks at the ST and NT. All OK.
Removed Partisol cassettes that ran on 9/26, 9/29, and 10/3.
Area along the Calument River near the ST-NW BAM has been paved with asphalt.
Checked the ambient temperature sensor on the NT-SW BAM at 13:05 CST. While
cleaning the sensor, I accidentally damaged the probe. Replaced the sensor with the
backup. Checked the new sensor output with the Delta Cal. NT-NW BAM = 15.2 deg C
Delta Cal = 15.1 deg C.
2014-10-07 | 08:50 G. Mazik
Performed QC Checks on the BAM monitors.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2014-10-08 | 08:45 G. Mazik
Checked on communications issue with the NT-SW BAM. Reset the data logger and
communications resumed.
2014-10-10 | 08:50 G. Mazik
Performed QC Checks on the Partisol monitors.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2014-10-13 | 08:55 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2014-10-16 | 09:10 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2014-10-21 | 09:40 G. Mazik
Conducted a site check of the monitoring equipment.
Performed QC Checks on the BAM monitors.
See specific tabs for additional information about that site.
2014-10-24 | 08:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2014-10-28 | 11:45 G. Mazik
Conducted a site check of the monitoring equipment.
Performed QC Checks on the Partisol monitors.
See specific tabs for additional information about that site.
2014-10-30 | 09:05 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2014-11-03 | 09:55 G. Mazik
Conducted a site check of the monitoring equipment.
Performed QC Checks on the BAM monitors.
See specific tabs for additional information about that site.
2014-11-05 | 11:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2014-11-10 | 10:35 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2014-11-11 | 09:00 G. Mazik
Checked NT-metals Partisol prior to project progress meeting with KCBX and URS.
2014-11-14 | 09:50 G. Mazik
Conducted a site check of the monitoring equipment.
Performed QC Checks on the Partisol monitors.
See specific tabs for additional information about that site.
2014-11-17 | 10:10 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2014-11-19 | 10:15 G. Mazik
Checking on communications issues with the NT-SW BAM monitor.
See specific tabs for additional information about that site.
2014-11-21 | 10:20 G. Mazik
Conducted a site check of the monitoring equipment.
Performed QC Checks on the BAM monitors.
See specific tabs for additional information about that site.
2014-11-26 | 08:35 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2014-12-01 | 10:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2014-12-05 | 09:40 G. Mazik
Conducted a site check of the monitoring equipment.
Performed QC Checks on the BAM monitors.
Performed QC Checks on the Partisol monitors.
See specific tabs for additional information about that site.
2014-12-06 | 07:15 G. Mazik
Onsite to investigate a problem with the NT-NW BAM monitor.
2014-12-08 | J. Sorkness conducted audits of the BAM and Partisol monitors.
Conducted site checks. OK.
2014-12-09 | J. Sorkness conducted audits of the meteorological systems.
2014-12-12 | 09:40 G. Mazik
Conducted site checks. OK.
See specific tabs for additional information about that site.
2014-12-16 | 10:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2014-12-18 | 09:45 G. Mazik
Performed QC Checks on the BAM monitors.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2014-12-23 | 09:50 G. Mazik
Conducted site checks. OK.
See specific tabs for additional information about that site.
2014-12-26 | 11:15 G. Mazik
Onsite to investigate a problem with the NT-Carbon Partisol monitor.
2014-12-30 | 09:10 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-01-02 | 09:30 G. Mazik
Performed QC Checks on the Partisol monitors.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-01-05 | 10:10 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-01-08 | 10:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-01-10 | 11:15 G. Mazik
Onsite to manually load Partisol cassettes and check on alarm conditions.
2015-01-12 | 11:10 G. Mazik
Performed QC Checks on the BAM monitors.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-01-13 | 09:45 G. Mazik
Began annual zeroing of the BAM monitors.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-01-16 | 09:30 G. Mazik
Concluded annual zeroing of the BAM monitors.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-01-20 | 10:45 G. Mazik
Input new zero values into the BAM monitors.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-01-23 | 09:45 G. Mazik
Performed QC checks on the Partisol samplers.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-01-26 | 10:15 G. Mazik
Performed QC checks on the BAM monitors.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-01-29 | 10:10 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-01-31 | 09:40 G. Mazik
Manually loaded filters into Partisol samplers for the 2/2 run date.
2015-02-04 | 10:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-02-06 | 10:15 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-02-10 | 10:00 G. Mazik
Performed QC checks on the BAM monitors.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-02-13 | 09:45 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-02-16 | 09:00 G. Mazik
Performed QC checks on the Partisol samplers.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-02-19 | 10:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-02-21 | 08:20 G. Mazik
Manually loaded filters into Partisol samplers for the 2/23 run date.
2015-02-24 | 09:20 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-02-27 | 10:45 G. Mazik
Performed QC checks on the BAM monitors.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-03-02 | 10:20 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-03-06 | 09:15 G. Mazik
Changed filter tapes on the BAM monitors.
Performed leak checks on the BAM monitors.
Cleaned PM10 heads on the BAM monitors.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-03-09 | 09:20 G. Mazik
Performed QC checks on the Partisol monitors.
Cleaned PM10 heads on the Partisol monitors.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-03-11 | 09:40 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-03-12 | 09:30 G. Mazik
Onsite to check on the operation of the Partisol shuttles at the NT.
2015-03-14 | 09:15 G. Mazik
Manually loaded filters into Partisol samplers for the 3/16 run date.
2015-03-16 | Todd advised that the electrician began rewiring the circuit to the ST-NE. He was supposed
to tell me but forgot. Supposedly, only a minute or two was lost during the change out.
3/16 was a Partisol run day. If we only lost a couple of minutes, the sample should be OK.
2015-03-17 | 09:00 G. Mazik
Planned to conduct QC checks on the BAM monitors, but was unable to do so due to
measured concentrations exceeding 100 ug/m3 on several units.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-03-19 | 08:30 G. Mazik, J. Sorkness
J. Sorkness completed audits of the ST and NT meteorological stations. OK.
J. Sorkness completed audits of the NT amd ST BAM monitors. OK.
See specific tabs for additional information about that site.
2015-03-20 | 06:15 G. Mazik, J. Sorkness
J. Sorkness completed audits of the ST and NT Partisol samplers. OK.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-03-24 | 09:10 G. Mazik
Performed QC checks on the BAM monitors. OK.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-03-27 | 09:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-03-31 | 09:00 G. Mazik
Performed leak checks on the BAM monitors. OK.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-04-01 | 08:30 G. Mazik, P. Hight
Reviewed equipment operation with P. Hight, who will fill in for me next week.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-04-08 | 09:00 G. Mazik
Performed QC checks on the BAM monitors. OK.
Performed QC checks on the Partisol monitors. OK.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-04-10 | 09:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-04-15 | 12:20 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-04-17 | 10:50 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-04-24 | 08:45 G. Mazik
Performed QC checks on the BAM monitors. OK.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-04-27 | 12:20 G. Mazik
Installed new alarm status wiring on the BAM monitors.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-04-30 | 10:50 G. Mazik
Changed BAM monitor filter tapes.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-05-04 | 09:00 G. Mazik
Performed QC checks on the Partisol samplers. OK.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-05-07 | 09:10 G. Mazik
Performed QC checks on the BAM monitors. OK.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-05-11 | 11:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-05-14 | 10:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-05-18 | 08:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-05-21 | 09:30 G. Mazik
Performed QC checks on the BAM monitors. OK.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-05-27 | 09:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-05-29 | 09:00 G. Mazik
Performed QC checks on the Partisol samplers. OK.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-06-01 | 08:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-06-04 | 08:40 G. Mazik
Performed QC checks on the BAM monitors. OK.
Cleaned BAM nozzles and screens.
Cleaned BAM PM10 heads.
Cleaned Partisol PM10 heads.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-06-09 | 08:45 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-06-12 | 09:00 G. Mazik
Performed QC checks on the Partisol samplers. OK.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-06-15 | 09:45 G. Mazik
Performed leak checks on the BAM monitors.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-06-16 | 09:45 G. Mazik
Replaced the temperature probe at the ST-Met station.
2015-06-17 | 07:45 G. Mazik, J. Sorkness
J. Sorkness performed audit checks on the BAM monitors.
J. Sorkness performed audit checks on the meteorological systems.
2015-06-18 | 07:00 G. Mazik, J. Sorkness
J. Sorkness performed audit checks on the Partisol samplers.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-06-22 | 08:50 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-06-24 | 07:10 G. Mazik
Changed out BAM filter tapes.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-06-30 | 07:45 G. Mazik
Performed QC checks on the Partisol samplers. OK.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-07-02 | 07:00 G. Mazik
Performed QC checks on the BAM monitors. OK.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-07-04 | High PM concentrations were measured across the network from 7/4/15 20:00 to
2015-07-05 | 7/5/15 04:00 CST. These concentrations were likely caused by neighborhood fireworks in
the area. Winds were calm during this period.
2015-07-07 | 09:15 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-07-10 | 08:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-07-14 | 09:00 G. Mazik
Performed QC checks on the BAM monitors. OK.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-07-16 | 09:15 G. Mazik
Performed QC checks on the Partisol samplers. OK.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-07-22 | 09:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-07-24 | 08:40 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-07-28 | 08:40 G. Mazik
Performed QC checks on the BAM monitors. OK.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-07-30 | 08:30 G. Mazik
Performed QC checks on the Partisol samplers. OK.
Power outage at the entire NT at 05:00 CST. Utility company repair techinicans indicated
the problem was caused by a bird shorting out the system.
Power was restored to all NT sites by 08:00 CST.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-08-05 | 08:55 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-08-07 | 08:15 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-08-11 | 09:20 G. Mazik
Performed BAM leak checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-08-12 | 07:15 G. Mazik, J. Sorkness
J. Sorkness performed audits of the BAM, Partisol, and meteorological systems.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-08-14 | 09:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-08-18 | A lightning strike at approximately 1900 CST caused a power failure at all ST sites except
ST-NE. The ST-NE site is connected to the office building, which has a backup generator.
2015-08-19 | 08:00 G. Mazik
Power still out at all ST monitoring sites except ST-NE.
Power restored to all ST monitoring sites by 16:30 CST.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-08-21 | 08:00 G. Mazik
Performed QC checks on the BAM monitors. OK.
Changed BAM filter tapes.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-08-28 | 08:45 G. Mazik
Took BAMs offline to record Last M and ABS readings.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-09-01 | 08:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-09-04 | 08:00 G. Mazik
Performed QC checks on the BAM monitors. OK.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-09-09 | 09:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-09-11 | 09:30 G. Mazik
Performed Partisol QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-09-15 | 08:30 G. Mazik, D. Szatkowski
D. Szatkowski cleaned the membrane on the NT-SE BAM.
Replaced the NT BAM pumps with rebuilt ones.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-09-16 | 07:30 G. Mazik, D. Szatkowski
Replaced the ST BAM pumps with rebuilt ones.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-09-22 | 08:05 G. Mazik
Performed QC checks on the BAM monitors. OK.
Begin 72-hour BAM zero calibrations.
Change zero offsets to 0.0000.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-09-25 | 09:00 G. Mazik
Concluded BAM zero calibrations.
Adjusted BAM zero offsets to former values.
New BAM zero offsets will be input next week.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-09-28 | 09:30 G. Mazik
Placed new zero offsets into all BAMs.
Started a 72-hour zero test on BAM SN J5560 located in the NT lunch room.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-09-30 | 09:25 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
Checked on the 72-hour zero test for BAM SN J5560.
2015-10-02 | 08:30 G. Mazik
Performed Partisol QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-10-05 | 08:55 G. Mazik
Performed QC checks on the BAM monitors. OK.
Replaced BAM SN M4792 at the NT-NW site with BAM SN J5560.
Electical power at the ST-NE site was transferred to ComEd emergency service.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-10-08 | 09:20 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-10-13 | 09:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-10-15 | 10:20 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-10-19 | 10:20 G. Mazik
Strong winds blowing from the south. All BAMs showing elevated concentrations.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-10-22 | 09:15 G. Mazik
Performed QC checks on the BAM monitors. OK.
Changed BAM filter tapes.
Performed Partisol QC checks. OK.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-10-23 | 10:00 G. Mazik
Onsite to investigate an intermittent "D" error reported by the NT-NW BAM.
2015-10-27 | 08:50 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-10-30 | 10:40 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-11-02 | 09:30 G. Mazik
Electrical service to the NT-NE and NT-SE sites was changed over to permanent sources
to reduce the chance of future power failures.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-11-04 | 10:45 G. Mazik
Performed Partisol QC checks.
Started a zero test on the BAM that returned from the vendor.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-11-10 | 10:00 G. Mazik
Performed Partisol QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-11-12 | 11:00 G. Mazik
Replaced BAM SN J5560 with SN M4792 at the NT-NW site.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-11-13 | 09:45 G. Mazik
Power being changed out at the NT-NE site. Power was off from 09:00 to 11:00 CST
The air quality samplers came back online once power was restored.
2015-11-16 | 10:40 G. Mazik
Performed BAM leak checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-11-18 | 09:00 G. Mazik, J. Sorkness
J. Sorkness performed audits of the BAM and Partisol monitors.
The meteorological systems were not audited at this time due to safety concerns
with the high measured winds.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-11-23 | 10:45 G. Mazik
WS and WD sensors at the NT frozen due to heavy snowfall from 15:00 CST on 11-21-15
to 07:00 CST on 11-23-15.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-11-25 | 09:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-11-30 | 10:30 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-12-03 | 10:00 G. Mazik, P. Hight
Performed audits of the ST and NT meteorological systems.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-12-08 | 10:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-12-10 | 11:45 G. Mazik
Performed Partisol QC checks.
Cleaned all PM10 heads.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-12-14 | 10:15 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-12-16 | 10:00 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-12-22 | 09:30 G. Mazik
Changed BAM filter tapes.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-12-25 | Power failure at all NT sites from 07:00 to 09:00 CST.
The NT meterological station seems to have run off the battery backup during the
power failure.
2015-12-29 | 09:10 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2015-12-31 | 09:10 G. Mazik
Performed Partisol QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-01-04 | 09:40 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-01-05 | 13:00
Air monitoring at the NT discontinued.
2016-01-06 | 10:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-01-11 | 10:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-01-14 | 10:15 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-01-19 | 10:00 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-01-21 | 09:45 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-01-25 | 10:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-01-27 | 11:30 G. Mazik
Performed Partisol QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-02-01 | 10:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-02-04 | 09:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-02-08 | 10:00 G. Mazik
Performed BAM QC checks.
Changed BAM filter tapes.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-02-11 | 11:45 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-02-16 | 10:00 G. Mazik
Performed Partisol QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-02-19 | 09:45 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-02-23 | 09:55 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-02-24 | 09:30 G. Mazik
Onsite to check out problems with the NT-N BAM communications.
2016-02-26 | 11:45 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-02-29 | 10:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-03-04 | 09:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-03-08 | 09:40 G. Mazik
Performed BAM QC checks.
Performed Partisol QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-03-11 | 11:05 G. Mazik
Cleaned BAM PM10 heads.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-03-14 | 09:45 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-03-17 | 09:35 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
Cleaned Partisol PM10 heads.
Cleaned enclosure air conditioners and changed filters.
2016-03-23 | 08:20 G. Mazik, J. Sorkness
J. Sorkness performed audits of the PM10 samplers, Partisol monitors, and met systems.
Results indicate all systems to be operating within allowable project limits.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-03-25 | 08:50 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-04-01 | 08:10 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-04-02 | Power failure caused by high winds at the ST-N site beginning at 11:00 CST.
Power failure caused by high winds at the ST-SW site beginning at 19:00 CST.
2016-04-03 | Power restored to the ST-SW site at 10:00 CST.
Power restored to the ST-N site at 18:00 CST.
2016-04-04 | 09:05 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-04-07 | 09:25 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-04-11 | 09:05 G. Mazik
Performed preventative maintennace on the BAMs.
Changed BAM filter tapes.
Performed Partisol QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-04-14 | 09:25 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-04-18 | 09:15 G. Mazik
Postponed planned BAM QC checks due to high concentrations on several monitors.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-04-21 | 09:20 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-04-25 | 09:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-04-29 | 08:30 G. Mazik
Cleaned BAM PM10 heads.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-05-03 | 09:00 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-05-05 | 08:50 G. Mazik
Performed Partisol QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-05-09 | 09:10 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-05-12 | 10:20 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-05-16 | 08:50 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-05-19 | 09:15 G. Mazik
Performed Partisol QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-05-24 | 09:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-05-25 | Elevated concentrations reported on all BAM monitors due to high winds.
2016-05-27 | 08:05 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-05-28 | Power failure at the ST-NE BAM site during hour 21:00.
2016-05-31 | 09:15 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-06-02 | 09:10 G. Mazik
Performed Partisol QC checks.
Cleaned BAM enclosure air conditioner filters.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-06-07 | 07:30 G. Mazik, J. Sorkness
J. Sorkness performed audits of the PM10 samplers, Partisol monitors, and met systems.
Results indicate all systems to be operating within allowable project limits, except for WD.
The WD sensor was changed out during the audit, after consulting with J. Clarke.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-06-10 | 06:30 G. Mazik
Change-out BAM filter tapes.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-06-14 | 12:50 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-06-15 | Picked up Partisol filters for preparation and transport to the laboratory for storage.
2016-06-16 | Hours 2100-2200 are missing from the KCBX RAQIS database.
2016-06-20 | 08:15 G. Mazik
Cleaned BAM PM10 heads.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-06-23 | 08:25 G. Mazik
Cleaned enclosure air conditioning filters.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-06-27 | 07:40 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-06-28 | Hours 2100-2200 are missing from the KCBX RAQIS database.
2016-06-30 | 08:40 G. Mazik
Performed Partisol QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
Hour 2100 is missing from the KCBX RAQIS database.
2016-07-06 | 09:40 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-07-08 | 08:20 G. Mazik
Cleaned BAM PM10 heads.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-07-11 | 09:20 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-07-12 | Hours 1900-2000 are missing from the KCBX RAQIS database.
2016-07-14 | 10:40 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-07-19 | 08:30 G. Mazik
Performed Partisol QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-07-22 | 09:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-07-26 | 08:50 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-07-28 | 11:00 G. Mazik
Changed out filter tapes.
Cleaned BAM nozzles and vanes.
Cleaned BAM capstan shafts and pinch roller tires.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-08-02 | 09:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-08-04 | Power failure at the ST-SW site during hour 20:00 CST.
2016-08-05 | 08:40 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-08-08 | Hours 0800-0900 are missing from the KCBX RAQIS database.
2016-08-09 | 08:35 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-08-12 | 09:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-08-15 | 09:00 G. Mazik
Performed Partisol QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-08-18 | 09:50 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-08-22 | 09:10 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-08-24 | Power failure at the ST-SW site during hour 19:00 CST.
2016-08-26 | 09:40 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
Cleaned BAM shelter air conditioner filters.
2016-08-29 | 09:05 G. Mazik
Performed Partisol QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-09-01 | 07:30 G. Mazik, J. Sorkness
J. Sorkness performed audits of the PM10 samplers, Partisol monitors, and met systems.
Results indicate all systems to be operating within allowable project limits.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-09-06 | 09:25 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-09-09 | 09:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
Cleaned BAM PM10 heads.
2016-09-13 | 09:10 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-09-16 | 09:05 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-09-19 | 09:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-09-22 | 08:40 G. Mazik
Changed out the BAM filter tapes.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-10-01 | 09:10 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-10-04 | 09:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-10-07 | 08:30 G. Mazik
Performed Partisol QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-10-11 | 10:10 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-10-14 | 09:05 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
Cleaned BAM and Partisol PM10 heads.
2016-10-17 | 09:20 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-10-20 | 08:50 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-10-25 | 08:50 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-10-28 | 08:55 G. Mazik
Performed Partisol QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-11-02 | 08:40 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-11-04 | 08:20 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
Cleaned BAM enclosure AC filters.
Cleaned Partisol rain hoods and air screens.
2016-11-08 | 09:45 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-11-11 | 09:20 G. Mazik
Performed BAM QC checks, except for ST-N. Due to the recording of a
concentration >100 ug/m3, the ST-N BAM could not be taken offline for testing.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-11-15 | 09:05 G. Mazik
Performed Partisol QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-11-18 | 09:40 G. Mazik
High southerly winds generating elevated concentrations on the ST-SW and
ST-SE monitors. Salt trucks being loaded on the property to the south.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-11-21 | Hours 1700 and 1800 CST recorded on the ST-N BAM were invalid due to heavy loading
at the filter tape, and subsequent pressure drop, during a fire in the building to the north
of the sampler, as reported by KCBX staff.
2016-11-22 | 09:45 G. Mazik
Performed BAM QC checks, except for ST-SW. Due to recording of a
concentration >100 ug/m3, the ST-SW BAM could not be taken offline for testing.
Changed out filter tapes, except for ST-SW.
Cleaned BAM nozzles and vanes, except for ST-SW.
Cleaned BAM capstan shafts and pinch roller tires, except for ST-SW.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-11-25 | 08:45 G. Mazik
Performed BAM QC checks on ST-SW.
Changed out filter tape on ST-SW.
Cleaned BAM nozzles and vanes on ST-SW.
Cleaned BAM capstan shafts and pinch roller tires on ST-SW.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-11-30 | 13:40 G. Mazik, J. Sorkness
J. Sorkness performed audits of the Partisol monitors, ST-NE BAM, and ST-N BAM.
Results indicate all systems to be operating within allowable project limits.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-12-02 | 08:15 G. Mazik, J. Sorkness
J. Sorkness performed audits of the ST-NW BAM, ST-SW BAM, ST-SE BAM and
meteorological systems.
Results indicate all systems to be operating within allowable project limits.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-12-06 | 10:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-12-09 | 09:25 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-12-13 | 09:55 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-12-16 | 08:15 G. Mazik
Performed Partisol QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-12-20 | 09:35 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-12-22 | 09:15 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
Cleaned BAM PM10 heads.
2016-12-28 | 12:20 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2016-12-30 | 10:05 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-01-03 | 15:20 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-01-06 | 13:50 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-01-09 | 10:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-01-13 | 10:00 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-01-17 | 10:15 G. Mazik
Changed out BAM filter tapes.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-01-23 | 09:55 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-01-27 | 09:30 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-01-31 | 10:20 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-02-06 | 09:50 G. Mazik
Performed final Partisol QC checks. The Partisol sampling program has been discontinued.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-02-13 | 10:20 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-02-17 | 09:20 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-02-23 | 10:05 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-02-27 | 09:20 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-03-01 | 16:00 G. Mazik
RAQIS has not been updating from the ST-SW BAM since 19:00 CST on 2/28/17.
A lightning storm had passed through the area that evening. Found the data logger
screen blank. Power cycled the data logger and NL201, without success. Tried the spare
data logger. It would not get past the initial screen. Sending both units to D. Szatkowski
for evaluation. The BAM is still collecting valid data.
2017-03-03 | 09:40 G. Mazik
Installed replacement data logger at the ST-SW site. Seems to be working properly.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-03-08 | 10:15 G. Mazik
Cleaned BAM PM10 heads.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-03-16 | 08:40 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-03-20 | 08:55 G. Mazik
Cleaned enclosure AC filters.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-03-24 | 07:00 G. Mazik, J. Sorkness
J. Sorkness performed audits of the BAMs and meteorological systems. All OK.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-03-27 | 09:00 G. Mazik
Changed out BAM filter tapes.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-03-31 | 08:20 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-04-07 | 08:35 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-04-10 | 14:45 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-04-13 | 08:55 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-04-18 | 15:55 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-04-21 | 11:20 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-04-25 | 09:05 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-04-28 | 09:25 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-05-05 | 08:40 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-05-12 | 08:40 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-05-16 | 09:25 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
Elevated concentrations reported on all BAM monitors due to high winds.
2017-05-17 | Elevated concentrations reported on all BAM monitors due to high winds.
2017-05-19 | 12:55 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
Cleaned BAM PM10 heads.
2017-05-23 | 09:50 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-05-26 | 07:45 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-05-30 | 09:10 G. Mazik
Changed out BAM filter tapes.
Cleaned BAM nozzles and vanes.
Cleaned BAM capstan shafts and pinch roller tires.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-06-05 | 08:55 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-06-08 | 09:00 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-06-13 | 09:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-06-14 | A thunder storm knocked out power to the ST-SE BAM and met station at 12:00 CST.
Due to the battery backup, the meteorological station operated until 17:00 CST.
2017-06-15 | The blown fuse was replaced by the local power company at 11:20 CST.
The ST-SE BAM and meteorological systems began recording valid data.
2017-06-20 | 09:15 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-06-23 | 08:45 G. Mazik
Performed BAM QC checks.
Cleaned shelter AC filters.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-06-29 | 06:45 G. Mazik, J. Sorkness
J. Sorkness performed audits of the BAMs and meteorological systems. All OK.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-07-05 | 11:15 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-07-11 | 08:00 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-07-14 | 08:25 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-07-18 | 09:05 G. Mazik
Cleaned BAM PM10 heads.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-07-21 | 07:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-07-24 | 07:30 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-07-27 | 14:35 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-07-31 | 09:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-08-03 | 12:00 G. Mazik
Planned to change out BAM filter tapes, but could not because two the the BAMs
were reading >100 ug/m3.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-08-04 | 11:45 G. Mazik
Changed out BAM filter tapes.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-08-08 | 09:05 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-08-09 | 14:00 G. Mazik
Onsite to investigate a power failure at the ST-SW BAM. One of the KCBX employees
accidentally turned off a light switch that supplies power to the BAM. The employee
turned the light back on and the BAM resumed normal operation.
2017-08-11 | 09:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-08-15 | 09:20 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-08-16 | 14:20 G. Mazik
Replaced the AC filters on the shelters.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-08-22 | 09:10 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-08-25 | 08:40 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-08-29 | 08:05 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-08-31 | 14:25 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-09-05 | 09:00 G. Mazik
Cleaned BAM PM10 heads.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-09-08 | 07:30 G. Mazik, J. Sorkness
J. Sorkness performed audits of the BAMs and meteorological systems. All OK.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-09-12 | 09:20 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-09-15 | 09:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-09-19 | 14:35 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-09-22 | 08:50 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-09-25 | 09:00 D. Hubbard, D. Szatkowski
Investigated the loss of communications with the ST-SW BAM. Found the battery backup
side of the UPS no longer holds a charge. Transferred communication equipment
power plugs to other side of the UPS without battery backup.
2017-10-02 | 09:20 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-10-05 | 09:40 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-10-09 | 09:20 G. Mazik
Changed out BAM filter tapes.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-10-12 | 09:15 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-10-16 | 09:40 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-10-20 | 08:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-10-24 | 09:45 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-10-27 | 09:00 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-11-06 | 10:40 G. Mazik
Rebuilt the BAM sample pumps on ST-SW and ST-SE.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-11-10 | 09:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-11-14 | 10:35 G. Mazik
Rebuilt the BAM sample pumps on ST-NW, ST-N, and ST-NE.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-11-17 | 08:05 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-11-24 | 09:20 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-11-27 | 10:20 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-11-30 | 07:15 G. Mazik, J. Sorkness
J. Sorkness performed audits of the BAMs and meteorological systems. All OK.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-12-05 | 10:45 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-12-12 | 10:35 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-12-15 | 10:05 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-12-19 | 10:00 G. Mazik
Changed out BAM filter tapes.
Cleaned BAM nozzles and vanes.
Cleaned BAM capstan shafts and pinch roller tires.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-12-22 | 09:05 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-12-26 | 10:30 G. Mazik
Meteorological system backup battery failed. Replaced.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2017-12-29 | 09:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-01-02 | The ST-N and ST-NE monitors recorded several flow errors in the morning, likely due to
extreme low temperatures. Both monitors resumed normal operation at 08:00 CST.
2018-01-03 | 10:15 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-01-08 | A power failure was recorded at the ST-NW site due to a blown fuse. Power to the ST-SW
monitor was turned off during the fuse replacement.
2018-01-09 | 10:20 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-01-10 | A power failure was recorded at the ST-NW site due to a blown fuse. Power to the ST-SW
monitor was turned off during the fuse replacement.
15:20 G. Mazik
Inspected the ST-NW and ST-SW BAMs following fuse replacement.
2018-01-12 | 10:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-01-17 | 10:20 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-01-19 | 09:40 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-01-22 | 11:05 G. Mazik
Cleaned enclosure AC filters.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-01-25 | 10:30 G. Mazik
Cleaned PM10 heads.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-01-29 | 10:20 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-02-02 | 10:40 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-02-06 | 11:20 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-02-13 | 10:45 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-02-16 | 10:05 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-02-20 | 11:00 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-02-23 | 07:55 G. Mazik
Changed out BAM filter tapes.
Cleaned BAM nozzles and vanes.
Cleaned BAM capstan shafts and pinch roller tires.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-02-26 | 10:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-03-01 | 10:50 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-03-06 | 10:10 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-03-09 | 09:50 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-03-13 | 09:15 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-03-16 | 09:40 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-03-20 | 09:25 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-03-23 | 09:10 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-03-27 | 09:35 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-03-29 | 06:45 G. Mazik, J. Sorkness
J. Sorkness performed audits of the BAMs and meteorological systems. All OK.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-04-03 | 09:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-04-06 | 09:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-04-10 | 09:15 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-04-13 | 08:55 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-04-16 | 11:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-04-19 | 09:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-04-23 | 09:20 G. Mazik
Performed BAM QC checks, except for ST-SE.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-04-26 | 09:25 G. Mazik
Performed BAM QC checks on ST-SE.
Changed out BAM filter tapes.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-05-01 | 09:15 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-05-04 | 09:20 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-05-11 | 12:15 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-05-18 | 08:45 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-05-22 | 09:50 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-05-25 | 08:30 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-05-29 | 09:10 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-06-01 | 08:45 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-06-05 | 09:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-06-08 | 09:45 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-06-11 | 13:30 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-06-14 | 08:45 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-06-18 | 08:55 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-06-21 | 09:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-06-25 | 09:15 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-06-28 | 07:30 G. Mazik, J. Sorkness
J. Sorkness performed audits of the BAMs and meteorological systems.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-07-02 | 08:45 G. Mazik
Changed out BAM filter tapes.
Cleaned BAM nozzles and vanes.
Cleaned BAM capstan shafts and pinch roller tires.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-07-05 | 10:00 G. Mazik
Cleaned BAM PM10 heads.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-07-10 | 08:40 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-07-13 | 08:40 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-07-16 | 09:10 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-07-19 | 09:15 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-07-24 | 09:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-07-27 | 09:20 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-07-30 | 09:15 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-08-03 | 08:10 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-08-07 | 09:15 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-08-10 | 09:00 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-08-14 | 09:00 G. Mazik
Cleaned enclosure AC filters.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-08-17 | 09:15 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-08-23 | 09:20 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-08-28 | 09:15 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-08-30 | 09:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-09-04 | 09:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-09-07 | 09:20 G. Mazik
Changed out BAM filter tapes.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-09-11 | 10:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-09-14 | 09:20 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-09-18 | 09:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-09-21 | 08:40 G. Mazik
Cleaned BAM PM10 heads.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-09-24 | 09:50 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-09-27 | 07:30 G. Mazik, J. Sorkness
J. Sorkness performed audits of the BAMs and meteorological systems, except for
PM10 flows.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-09-28 | 08:40 G. Mazik
Performed PM10 flow audit checks with DeltaCal s/n 1301.
2018-10-02 | 10:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-10-05 | 10:45 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-10-12 | 08:05 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-10-16 | 09:35 G. Mazik
Replaced WD sensor.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-10-19 | 08:25 G. Mazik
Repaired a loose power wire to the data logger.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-10-23 | 09:25 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-10-26 | 08:15 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-10-30 | 09:00 G. Mazik
Cleaned BAM enclosure AC filters.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-11-02 | 08:40 G. Mazik
Changed out the filter tape on the ST-SE BAM.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-11-06 | 13:20 G. Mazik
RAQIS stopped recording met data. Power cycled the data logger and data
collection resumed. No data lost.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-11-12 | 09:55 G. Mazik
Changed out the filter tapes on the ST-NW, ST-N, ST-SW, and ST-NE BAMs.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-11-16 | 09:55 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-11-27 | 10:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-11-30 | 10:00 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-12-03 | 10:40 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-12-05 | 07:30 G. Mazik
Onsite to investigate a tape error message on ST-NW on 12/4 at 11:00.
Found the tape to be in good shape and the BAM to be operating normally.
2018-12-06 | 08:15 G. Mazik, J. Sorkness
J. Sorkness performed audits of the BAMs and meteorological systems.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-12-11 | 10:45 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-12-14 | 10:30 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-12-18 | 10:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-12-21 | 10:05 G. Mazik
Cleaned BAM PM10 heads.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-12-28 | 09:45 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2018-12-31 | 09:05 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-01-08 | 10:40 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-01-11 | 10:00 G. Mazik
Performed BAM QC checks.
Changed out the filter tape on the ST-SE BAM.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-01-15 | 10:10 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-01-18 | 10:00 G. Mazik
Changed out the filter tapes on the ST-NW, ST-N, ST-SW, and ST-NE BAMs.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-01-24 | 10:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-01-25 | 10:50 G. Mazik
Replaced the met system charger and battery due to a failure at 06:00 CST.
2019-01-30 | Extreme low ambient temperatures are causing flow errors on all BAMs.
D. Szatkowski contacted the vendor, who indicated the recorded PM10 data are still
valid.
2019-01-31 | Extreme low ambient temperatures are causing flow errors on all BAMs.
2019-02-01 | 10:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-02-07 | 12:50 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-02-11 | 10:10 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-02-15 | 10:10 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-02-18 | 12:30 G. Mazik
Replaced the NL201 on the met system.
2019-02-21 | 13:35 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-02-25 | 10:20 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-03-01 | 10:15 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-03-06 | 11:30 G. Mazik
Performed BAM QC checks.
Changed out the meteorological system data logger.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-03-12 | 09:25 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-03-15 | 09:25 G. Mazik
Trimmed the ends of the BAM sample pump lines.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-03-18 | 09:30 G. Mazik
Performed leak checks on the BAM monitors.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-03-20 | 07:40 G. Mazik, M. Wolfskill
M. Wolfskill performed audits of the BAMs and meteorological systems.
Changed out the filter tape on the ST-SE BAM.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-03-25 | 09:05 G. Mazik
Performed BAM QC checks.
Changed out the filter tapes on the ST-NW, ST-N, ST-SW, and ST-NE BAMs.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-03-29 | 10:15 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-04-02 | 09:30 G. Mazik
Cleaned BAM PM10 heads.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-04-05 | 09:35 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-04-09 | 09:20 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
High concentrations measured during hours 00:00 and 20:00, likely the result of fugitive
dust from properties to the north. Winds were from the N to NNE.
2019-04-12 | 09:15 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-04-16 | 08:50 G. Mazik
Inspected enclosure HVAC filters. Found all clean.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-04-18 | 09:05 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-04-23 | 09:55 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-04-26 | 09:15 G. Mazik
Planned to perform BAM QC checks, but could not due to a flow issue with the DeltaCal.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-04-30 | 09:10 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-05-03 | 09:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-05-06 | 10:35 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-05-10 | 09:20 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-05-16 | 09:10 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-05-22 | 12:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
Elevated concentrations measured at all sites during hours 13:00-16:00 CST.
Winds were from the south during this period.
2019-05-24 | 09:05 G. Mazik
Changed out the filter tape on the ST-SE BAM.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-05-28 | 09:50 G. Mazik
Performed BAM QC checks.
Changed out the filter tapes on the ST-NW, ST-N, ST-SW, and ST-NE BAMs.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-06-04 | 09:15 G. Mazik
Cleaned PM10 sampler heads.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-06-05 | 11:00 G. Mazik
Replaced the charger and battery on the meteorological system. It had failed at 13:00 CST
on 6/4/2019.
2019-06-07 | The ST-NE and ST-SE sites had a power failure at 14:00 CST.
2019-06-11 | 09:20 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-06-12 | Elevated concentrations measured at all sites during hours 07:00-09:00 CST.
Winds were from the south during this period.
2019-06-14 | 09:15 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-06-18 | 09:25 G. Mazik
Cleaned shelter HVAC filters.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-06-19 | Elevated concentrations measured at all sites during hour 05:00 CST.
Winds were from the northeast during this period.
2019-06-21 | 09:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-06-26 | 07:00 G. Mazik, M. Wolfskill
M. Wolfskill performed audits of the BAMs and meteorological systems.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-06-27 | Elevated concentrations measured at all sites during hour 18:00 CST.
Winds were from the north during this period.
2019-06-28 | 08:30 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-07-02 | 08:45 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-07-05 | 08:30 G. Mazik
Cleaned shelter HVAC filters.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-07-09 | 08:55 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-07-12 | 08:40 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-07-16 | 09:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-07-19 | 08:15 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-07-23 | 09:05 G. Mazik
Cleaned PM10 sampler heads.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-07-26 | 09:00 G. Mazik
Performed BAM QC checks.
Changed out the filter tape on the ST-SE BAM.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-07-30 | 08:50 G. Mazik
Changed out the filter tapes on the ST-NW, ST-N, ST-SW, and ST-NE BAMs.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-08-02 | 09:10 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-08-06 | 08:50 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-08-09 | 09:30 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-08-13 | 09:00 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-08-16 | 09:10 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-08-20 | 09:10 G. Mazik
Cleaned shelter HVAC filters.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-08-23 | 09:05 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-08-27 | 09:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-09-03 | 09:55 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-09-05 | 10:10 G. Mazik
Checked on communications issues with the ST-SE shelter. B. Stringer is sending me
a new router.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-09-06 | 10:00 G. Mazik
Installed the new router. B. Stringer checked the communication with the ST-SE shelter
and found everything to be back online. No data were lost.
2019-09-10 | 09:25 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-09-11 | All sites experienced a power failure during hours 03:00 and 04:00 CST except for ST-NE.
2019-09-13 | 09:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-09-17 | 09:30 G. Mazik
Replaced the Lantronix at the ST-SE shelter due to intermittent communications issues.
2019-09-19 | 07:30 G. Mazik, M. Wolfskill
M. Wolfskill performed audits of the BAMs and meteorological systems.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-09-20 | 09:30 G. Mazik
Inspected the meteorological system after reviewing the wind data and finding hours 03:00
and 04:00 recording all zeros or missing. Everything seems to be operating normally.
Checked with B. Stringer, and he indicated the wind instruments look good from his end.
The area experienced very heavy rains during that period.
2019-09-24 | 09:30 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-09-27 | 09:10 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-09-30 | 09:30 G. Mazik
Changed out the filter tape on the ST-SE BAM.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-10-04 | 08:30 G. Mazik
Changed out the filter tapes on the ST-NW, ST-N, and ST-SW BAMs.
Replaced the ST-NE BAM with the spare.
Performed QC checks on the new ST-NE BAM.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-10-07 | 09:45 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-10-09 | 10:05 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-10-15 | 09:15 G. Mazik
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-10-16 | 09:10 G. Mazik
Checked on ST-NW communications.
2019-10-21 | 09:20 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-10-24 | 09:25 G. Mazik
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-10-28 | 09:45 G. Mazik
Corrected the filter tape issue on the ST-NW BAM.
Performed BAM QC checks.
Conducted a site check of the monitoring equipment.
See specific tabs for additional information about that site.
2019-10-29 | 09:10 G. Mazik
Checked on ST-NW communications.

## ST-NW
2014-10-07 | Performed BAM QC checks starting at 12:20 CST.
Cleaned PM10 sample head.
Leak check as found was 0.3, so I did not clean the nozzle.
Conducted site checks. OK.
2014-10-10 | Conducted site checks. OK.
2014-10-13 | Conducted site checks. OK.
2014-10-16 | Conducted site checks. OK.
2014-10-21 | Performed BAM QC checks starting at 09:10 CST.
Conducted site checks. OK.
2014-10-24 | Conducted site checks. OK.
2014-10-28 | Conducted site checks. OK.
2014-10-30 | Conducted site checks. OK.
2014-11-03 | Performed BAM QC checks starting at 13:00 CST.
Conducted site checks. OK.
2014-11-05 | Conducted site checks. OK.
2014-11-10 | Conducted site checks. OK.
2014-11-14 | Conducted site checks. OK.
2014-11-17 | Conducted site checks. OK.
2014-11-21 | Performed BAM QC checks starting at 10:40 CST.
Replaced filter tape.
Conducted site checks. OK.
2014-11-26 | Conducted site checks. OK.
Changed time from 09:49:10 to 09:48 CST.
2014-12-01 | Conducted site checks. OK.
2014-12-05 | Performed BAM QC checks starting at 10:03 CST.
Conducted site checks. OK.
Cleaned PM10 sample head.
2014-12-08 | J. Sorkness conducted audits of the BAM.
Conducted site checks. OK.
2014-12-12 | Conducted site checks. OK.
2014-12-16 | Conducted site checks. OK.
2014-12-18 | Performed BAM QC checks starting at 10:15 CST.
Conducted site checks. OK.
2014-12-23 | Conducted site checks. OK.
2014-12-30 | Conducted site checks. OK.
2015-01-02 | Conducted site checks. OK.
2015-01-05 | Conducted site checks. OK.
2015-01-08 | Conducted site checks. OK.
2015-01-12 | Perform QC checks beginning at 11:20 CST.
Conducted site checks. OK.
2015-01-13 | Begin annual zeroing at 09:45 CST.
Conducted site checks. OK.
Changed units' background concentration to 0.0000.
2015-01-16 | Conducted site checks. OK.
Concluded annual zeroing at 11:05 CST.
Changed units' background concentration to -0.0040.
2015-01-20 | Input new zero background concentration of -0.0064 into BAM at 10:30 CST.
Conducted site checks. OK.
2015-01-23 | Conducted site checks. OK.
2015-01-26 | Perform QC checks beginning at 10:30 CST.
Conducted site checks. OK.
2015-01-29 | Conducted site checks. OK.
Changed time from 10:50:10 to 10:49 CST.
2015-02-04 | Conducted site checks. OK.
2015-02-06 | Conducted site checks. OK.
2015-02-10 | Perform QC checks beginning at 12:10 CST.
Conducted site checks. OK.
2015-02-13 | Conducted site checks. OK.
2015-02-16 | Conducted site checks. OK.
2015-02-19 | Conducted site checks. OK.
2015-02-24 | Conducted site checks. OK.
2015-02-27 | Perform QC checks beginning at 12:15 CST.
Conducted site checks. OK.
2015-03-02 | Conducted site checks. OK.
2015-03-06 | Changed filter tape beginning at 12:50 CST.
Performed leak check = 0.3.
Cleaned PM10 head.
Conducted site checks. OK.
2015-03-09 | Conducted site checks. OK.
2015-03-11 | Conducted site checks. OK.
2015-03-17 | Conducted site checks. OK.
2015-03-19 | J. Sorkness performed audits beginning at 12:02 CST.
2015-03-20 | Conducted site checks. OK.
2015-03-24 | Perform QC checks beginning at 10:25 CST.
Conducted site checks. OK.
2015-03-27 | Conducted site checks. OK.
2015-03-31 | Perform leak checks beginning at 11:00 CST.
Leak as found = 0.6, as left = 0.5.
Conducted site checks. OK.
2015-04-01 | Conducted site checks. OK.
2015-04-08 | Perform QC checks beginning at 09:35 CST.
Conducted site checks. OK.
2015-04-10 | Conducted site checks. OK.
2015-04-15 | Conducted site checks. OK.
2015-04-17 | Conducted site checks. OK.
2015-04-24 | Perform QC checks beginning at 10:25 CST.
Conducted site checks. OK.
2015-04-27 | Conducted site checks. OK.
Installed new alarm wiring beginning at 11:40 CST.
2015-04-30 | Conducted site checks. OK.
Changed filter tape beginning at 11:20 CST.
2015-05-04 | Perform QC checks beginning at 10:15 CST.
Conducted site checks. OK.
2015-05-07 | Conducted site checks. OK.
2015-05-11 | Conducted site checks. OK.
2015-05-14 | Conducted site checks. OK.
2015-05-18 | Conducted site checks. OK.
2015-05-21 | Perform QC checks beginning at 11:05 CST.
Conducted site checks. OK.
2015-05-27 | Conducted site checks. OK.
2015-05-29 | Conducted site checks. OK.
2015-06-01 | Conducted site checks. OK.
2015-06-04 | Perform QC checks beginning at 11:30 CST.
Cleaned nozzle and screen.
Cleaned PM10 head.
Conducted site checks. OK.
2015-06-09 | Conducted site checks. OK.
2015-06-12 | Conducted site checks. OK.
2015-06-15 | Performed leak check beginning at 12:05 CST.
Leak check = 0.3.
Conducted site checks. OK.
2015-06-17 | J. Sorkness performed audit checks beginning at 08:40 CST.
2015-06-18 | Conducted site checks. OK.
2015-06-22 | Conducted site checks. OK.
Changed time from 09:47:05 to 09:46 CST.
2015-06-24 | Conducted site checks. OK.
Changed filter tape at 07:30 CST.
2015-06-30 | Conducted site checks. OK.
2015-07-02 | Performed QC checks beginning at 11:15 CST.
Conducted site checks. OK.
2015-07-07 | Conducted site checks. OK.
2015-07-10 | Conducted site checks. OK.
2015-07-14 | Performed QC checks beginning at 09:05 CST.
Conducted site checks. OK.
2015-07-16 | Conducted site checks. OK.
2015-07-22 | Conducted site checks. OK.
2015-07-24 | Conducted site checks. OK.
2015-07-28 | Performed QC checks beginning at 10:20 CST.
Conducted site checks. OK.
2015-07-30 | Conducted site checks. OK.
2015-08-05 | Conducted site checks. OK.
2015-08-07 | Conducted site checks. OK.
2015-08-11 | Performed a BAM leak check beginning at 10:50 CST.
Conducted site checks. OK.
2015-08-12 | J. Sorkness performed a quarterly audit beginning at 07:15 CST.
Conducted site checks. OK.
2015-08-14 | Conducted site checks. OK.
2015-08-18 | Power failure caused by a lightning strike at 19:00 CST.
2015-08-19 | Power restored. Valid data collection resumed at 16:00 CST.
2015-08-21 | Performed QC checks beginning at 11:15 CST.
Changed BAM filter tape.
Conducted site checks. OK.
2015-08-28 | Conducted site checks. OK.
Took BAM offline to record Last M and ABS readings at 10:05 CST.
Last M = 0.769, ABS = 0.750
2015-09-01 | Conducted site checks. OK.
2015-09-04 | Performed QC checks beginning at 09:05 CST.
Changed time from 09:05:05 to 09:04 CST.
Performed flow calibration. As left flows: 15.0=15.15, 18.4=18.40, and 16.7=16.68.
Conducted site checks. OK.
2015-09-09 | Conducted site checks. OK.
2015-09-11 | Conducted site checks. OK.
Cleaned PM10 head.
2015-09-15 | Conducted site checks. OK.
2015-09-16 | Replaced the sample pump with a rebuilt one. Hours 09:00 - 11:00 CST are invalid due to
the change out.
Conducted site checks. OK.
2015-09-22 | Performed QC checks beginning at 09:30 CST.
Began 72-hour zero calibration following QC checks.
Changed zero offset to 0.0000.
Conducted site checks. OK.
2015-09-25 | Concluded zero calibration.
Adjusted zero offset to former number of -0.0064.
New zero offset numbers will be input next week.
Valid data collection resumed at 11:00 CST.
Conducted site checks. OK.
2015-09-28 | Placed new zero offset into monitor at 10:20 CST.
The new zero offset is -0.0022.
Conducted site checks. OK.
2015-09-30 | Conducted site checks. OK.
2015-10-02 | Conducted site checks. OK.
2015-10-05 | Performed QC checks beginning at 11:50 CST.
Conducted site checks. OK.
2015-10-06 | KCBX power failure from 02:00 to 03:00 CST.
2015-10-08 | Conducted site checks. OK.
2015-10-13 | Conducted site checks. OK.
2015-10-15 | Conducted site checks. OK.
2015-10-19 | Conducted site checks. OK.
2015-10-22 | Performed QC checks beginning at 12:05 CST.
Changed filter tape.
Conducted site checks. OK.
2015-10-27 | Conducted site checks. OK.
2015-10-29 | Conducted site checks. OK.
2015-11-02 | Conducted site checks. OK.
Changed time from 12:35:00 to 12:36 CST.
2015-11-04 | Performed QC checks beginning at 10:40 CST.
Conducted site checks. OK.
2015-11-10 | Conducted site checks. OK.
2015-11-12 | Conducted site checks. OK.
2015-11-16 | Performed a leak check beginning at 12:35 CST.
Leak = 0.4.
Conducted site checks. OK.
2015-11-18 | J. Sorkness performed a quarterly audit beginning at 09:00 CST.
Conducted site checks. OK.
2015-11-23 | Conducted site checks. OK.
2015-11-25 | Conducted site checks. OK.
2015-11-30 | Performed QC checks beginning at 12:40 CST.
Conducted site checks. OK.
2015-12-03 | Conducted site checks. OK.
2015-12-08 | Conducted site checks. OK.
2015-12-10 | Conducted site checks. OK.
Cleaned PM10 head.
2015-12-14 | Conducted site checks. OK.
2015-12-16 | Performed QC checks beginning at 11:40 CST.
Conducted site checks. OK.
2015-12-22 | Changed filter tape beginning at 11:10 CST.
Conducted site checks. OK.
2015-12-29 | Performed QC checks beginning at 10:30 CST.
Conducted site checks. OK.
2015-12-31 | Conducted site checks. OK.
2016-01-04 | Conducted site checks. OK.
2016-01-06 | Conducted site checks. OK.
2016-01-11 | Conducted site checks. OK.
2016-01-14 | Conducted site checks. OK.
2016-01-19 | Performed QC checks beginning at 10:50 CST.
Conducted site checks. OK.
2016-01-21 | Conducted site checks. OK.
2016-01-25 | Conducted site checks. OK.
2016-01-27 | Conducted site checks. OK.
2016-02-01 | Conducted site checks. OK.
2016-02-04 | Conducted site checks. OK.
2016-02-08 | Performed QC checks beginning at 09:35 CST.
Changed filter tape.
Conducted site checks. OK.
2016-02-11 | Conducted site checks. OK.
2016-02-16 | Conducted site checks. OK.
2016-02-19 | Conducted site checks. OK.
2016-02-23 | Performed QC checks beginning at 10:05 CST.
Conducted site checks. OK.
2016-02-26 | Conducted site checks. OK.
2016-02-29 | Conducted site checks. OK.
2016-03-04 | Conducted site checks. OK.
2016-03-08 | Performed QC checks beginning at 09:50 CST.
Changed time from 09:49:05 to 09:48 CST.
Conducted site checks. OK.
2016-03-11 | Conducted site checks. OK.
Cleaned PM10 head.
2016-03-14 | Conducted site checks. OK.
2016-03-17 | Conducted site checks. OK.
Cleaned enclosure air conditioner and changed filter.
2016-03-23 | J. Sorkness performed a quarterly audit beginning at 10:10 CST.
Conducted site checks. OK.
2016-03-25 | Conducted site checks. OK.
2016-04-01 | Conducted site checks. OK.
2016-04-03 | Power failure at 14:00 CST.
2016-04-04 | Performed QC checks beginning at 09:20 CST.
Conducted site checks. OK.
2016-04-07 | Conducted site checks. OK.
2016-04-11 | Performed preventative maintenace beginning at 09:50 CST.
Cleaned nozzle, screen, and tape rollers.
Changed filter tape.
Conducted site checks. OK.
2016-04-14 | Conducted site checks. OK.
2016-04-18 | Conducted site checks. OK.
2016-04-21 | Performed QC checks beginning at 09:35 CST.
Conducted site checks. OK.
2016-04-25 | Conducted site checks. OK.
2016-04-29 | Conducted site checks. OK.
Cleaned PM10 head.
2016-05-03 | Performed QC checks beginning at 09:00 CST.
Conducted site checks. OK.
Last M = 0.773, ABS = 0.750.
2016-05-05 | Conducted site checks. OK.
2016-05-09 | Conducted site checks. OK.
2016-05-12 | Conducted site checks. OK.
Changed time from 10:36:10 to 10:35 CST.
2016-05-16 | Performed QC checks beginning at 09:10 CST.
Conducted site checks. OK.
2016-05-19 | Conducted site checks. OK.
2016-05-24 | Conducted site checks. OK.
2016-05-25 | Elevated concentrations recorded due to high winds.
2016-05-27 | Performed QC checks beginning at 08:30 CST.
Conducted site checks. OK.
2016-05-31 | Conducted site checks. OK.
2016-06-02 | Conducted site checks. OK.
Cleaned BAM enclosure air conditioner filter.
2016-06-07 | J. Sorkness performed a quarterly audit beginning at 07:35 CST.
Conducted site checks. OK.
2016-06-10 | Conducted site checks. OK.
Changed out BAM filter tape at 06:50 CST,
2016-06-14 | Performed QC checks beginning at 13:10 CST.
Conducted site checks. OK.
2016-06-20 | Cleaned PM10 head.
Conducted site checks. OK.
2016-06-23 | Cleaned enclosure air conditioning filter.
Conducted site checks. OK.
2016-06-27 | Performed QC checks beginning at 08:10 CST.
Conducted site checks. OK.
2016-06-30 | Conducted site checks. OK.
2016-07-06 | Conducted site checks. OK.
2016-07-08 | Cleaned PM10 head.
Conducted site checks. OK.
2016-07-11 | Performed QC checks beginning at 09:35 CST.
Conducted site checks. OK.
2016-07-14 | Conducted site checks. OK.
2016-07-19 | Conducted site checks. OK.
2016-07-22 | Conducted site checks. OK.
2016-07-26 | Performed QC checks beginning at 08:50 CST.
Conducted site checks. OK.
2016-07-28 | Changed out filter tape beginning at 11:15 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
Conducted site checks. OK.
2016-08-02 | Conducted site checks. OK.
Changed time from 09:15:10 to 09:14 CST.
2016-08-05 | Conducted site checks. OK.
2016-08-09 | Performed QC checks beginning at 09:45 CST.
Conducted site checks. OK.
2016-08-12 | Conducted site checks. OK.
2016-08-15 | Conducted site checks. OK.
2016-08-18 | Conducted site checks. OK.
2016-08-22 | Performed QC checks beginning at 09:25 CST.
Conducted site checks. OK.
2016-08-26 | Conducted site checks. OK.
Cleaned shelter air condtioner filter.
Gravel has been has been installed to the north of the shelter.
Contractors pouring cement at several locations to the south of the shelter.
2016-08-29 | Conducted site checks. OK.
2016-09-01 | J. Sorkness performed a quarterly audit beginning at 07:45 CST.
Conducted site checks. OK.
2016-09-06 | Conducted site checks. OK.
2016-09-09 | Conducted site checks. OK.
Cleaned PM10 head.
2016-09-13 | Performed QC checks beginning at 09:20 CST.
Conducted site checks. OK.
2016-09-16 | Conducted site checks. OK.
2016-09-19 | Conducted site checks. OK.
KCBX was testing a new loading method to the south of the monitor.
2016-09-22 | Conducted site checks. OK.
Changed out filter tape starting at 09:10 CST.
2016-10-01 | Conducted site checks. OK.
2016-10-04 | Conducted site checks. OK.
2016-10-07 | Conducted site checks. OK.
2016-10-11 | Performed QC checks beginning at 10:15 CST.
Changed time from 10:23:05 to 10:22 CST.
Conducted site checks. OK.
2016-10-14 | Conducted site checks. OK.
Cleaned PM10 sample head.
2016-10-17 | Conducted site checks. OK.
2016-10-20 | Conducted site checks. OK.
2016-10-25 | Performed QC checks beginning at 09:10 CST.
Conducted site checks. OK.
Last M = 0.776, ABS = 0.750.
2016-10-28 | Conducted site checks. OK.
2016-11-02 | Conducted site checks. OK.
2016-11-04 | Conducted site checks. OK.
Cleaned enclosure AC filter.
2016-11-08 | Conducted site checks. OK.
2016-11-11 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2016-11-15 | Conducted site checks. OK.
2016-11-18 | Conducted site checks. OK.
2016-11-22 | Performed QC checks beginning at 10:10 CST.
Changed out filter tape.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
Conducted site checks. OK.
2016-11-25 | Conducted site checks. OK.
2016-11-30 | Conducted site checks. OK.
2016-12-02 | J. Sorkness performed a quarterly audit beginning at 09:50 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2016-12-06 | Conducted site checks. OK.
2016-12-09 | Conducted site checks. OK.
2016-12-13 | Performed QC checks beginning at 10:15 CST.
Conducted site checks. OK.
2016-12-16 | Conducted site checks. OK.
2016-12-20 | Conducted site checks. OK.
2016-12-22 | Conducted site checks. OK.
Cleaned PM10 head.
Changed time from 09:30:05 to 09:29 CST.
2016-12-28 | Performed QC checks beginning at 12:35 CST.
Conducted site checks. OK.
2016-12-30 | Conducted site checks. OK.
2017-01-03 | Conducted site checks. OK.
2017-01-06 | Conducted site checks. OK.
2017-01-09 | Conducted site checks. OK.
2017-01-13 | Performed QC checks beginning at 10:15 CST.
Conducted site checks. OK.
2017-01-17 | Performed filter tape change out beginning at 10:30 CST.
Conducted site checks. OK.
2017-01-23 | Conducted site checks. OK.
2017-01-27 | Performed QC checks beginning at 10:15 CST.
Conducted site checks. OK.
2017-01-31 | Conducted site checks. OK.
2017-02-06 | Conducted site checks. OK.
2017-02-13 | Performed QC checks beginning at 10:25 CST.
Conducted site checks. OK.
2017-02-17 | Conducted site checks. OK.
2017-02-23 | Conducted site checks. OK.
Changed time from 10:18:05 to 10:17 CST.
2017-02-27 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
2017-03-03 | Conducted site checks. OK.
2017-03-08 | Cleaned PM10 heads.
Conducted site checks. OK.
2017-03-16 | Performed QC checks beginning at 08:45 CST.
Conducted site checks. OK.
2017-03-20 | Cleaned enclosure AC filters.
Conducted site checks. OK.
2017-03-24 | J. Sorkness performed a quarterly audit beginning at 09:50 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2017-03-27 | Changed out BAM filter tape beginning at 09:10 CST.
Conducted site checks. OK.
2017-03-31 | Performed QC checks beginning at 09:20 CST.
Conducted site checks. OK.
2017-04-07 | Conducted site checks. OK.
2017-04-10 | Conducted site checks. OK.
2017-04-13 | Performed QC checks beginning at 09:10 CST.
Conducted site checks. OK.
2017-04-18 | Conducted site checks. OK.
2017-04-21 | Conducted site checks. OK.
2017-04-25 | Performed QC checks beginning at 09:40 CST.
Changed time from 09:40:05 to 09:39 CST.
Conducted site checks. OK.
2017-04-28 | Conducted site checks. OK.
2017-05-05 | Conducted site checks. OK.
2017-05-12 | Performed QC checks beginning at 08:50 CST.
Conducted site checks. OK.
2017-05-16 | Conducted site checks. OK.
2017-05-19 | Conducted site checks. OK.
Cleaned PM10 heads.
2017-05-23 | Conducted site checks. OK.
2017-05-26 | Performed QC checks beginning at 08:20 CST.
Conducted site checks. OK.
Last M = 0.773, ABS = 0.750.
2017-05-30 | Changed out filter tape beginning at 09:25 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
Leak check as found = 0.3. Leak check as left = 0.4
Conducted site checks. OK.
2017-06-05 | Conducted site checks. OK.
2017-06-08 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
2017-06-13 | Conducted site checks. OK.
2017-06-20 | Conducted site checks. OK.
2017-06-23 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
Cleaned shelter AC filter.
2017-06-29 | J. Sorkness performed a quarterly audit beginning at 08:30 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
Changed time from 07:34:20 to 07:33 CST.
2017-07-05 | Conducted site checks. OK.
2017-07-11 | Performed QC checks beginning at 09:15 CST.
Conducted site checks. OK.
2017-07-14 | Conducted site checks. OK.
2017-07-18 | Conducted site checks. OK.
Cleaned PM10 head.
2017-07-21 | Conducted site checks. OK.
2017-07-24 | Performed QC checks beginning at 08:45 CST.
Conducted site checks. OK.
2017-07-27 | Conducted site checks. OK.
2017-07-31 | Conducted site checks. OK.
2017-08-04 | Conducted site checks. OK.
Changed out BAM filter tape at 11:45 CST.
2017-08-08 | Performed QC checks beginning at 09:15 CST.
Conducted site checks. OK.
2017-08-11 | Conducted site checks. OK.
2017-08-15 | Conducted site checks. OK.
2017-08-16 | Conducted site checks. OK.
Replaced shelter AC filter.
2017-08-22 | Conducted site checks. OK.
2017-08-25 | Performed QC checks beginning at 09:05 CST.
Conducted site checks. OK.
2017-08-29 | Conducted site checks. OK.
2017-08-31 | Conducted site checks. OK.
2017-09-05 | Conducted site checks. OK.
Cleaned PM10 head.
2017-09-08 | J. Sorkness performed a quarterly audit beginning at 09:30 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2017-09-12 | Performed QC checks beginning at 09:25 CST.
Conducted site checks. OK.
2017-09-15 | Conducted site checks. OK.
2017-09-19 | Conducted site checks. OK.
2017-09-22 | Performed QC checks beginning at 08:05 CST.
Conducted site checks. OK.
2017-10-02 | Conducted site checks. OK.
2017-10-05 | Performed QC checks beginning at 10:10 CST.
Conducted site checks. OK.
2017-10-09 | Changed out BAM filter tape at 10:45.
Conducted site checks. OK.
2017-10-12 | Conducted site checks. OK.
2017-10-16 | Conducted site checks. OK.
2017-10-20 | Conducted site checks. OK.
2017-10-24 | Conducted site checks. OK.
2017-10-27 | Performed QC checks beginning at 09:20 CST.
Conducted site checks. OK.
2017-11-06 | Conducted site checks. OK.
2017-11-10 | Conducted site checks. OK.
2017-11-14 | Rebuilt the BAM sample pump. Offline at 11:10 CST for the change out.
Conducted site checks. OK.
2017-11-17 | Performed QC checks beginning at 08:15 CST.
Conducted site checks. OK.
2017-11-24 | Conducted site checks. OK.
2017-11-27 | Performed QC checks beginning at 10:30 CST.
Changed time from 10:32:05 to 10:31 CST.
Conducted site checks. OK.
2017-11-30 | J. Sorkness performed a quarterly audit beginning at 07:40 CST.
Results indicate all systems to be operating within allowable project limits.
Adjusted temperature from 4.9 to 6.4 degrees C.
Conducted site checks. OK.
2017-12-05 | Conducted site checks. OK.
2017-12-12 | Conducted site checks. OK.
2017-12-15 | Performed QC checks beginning at 10:20 CST.
Changed time from 10:21:10 to 10:20 CST.
Conducted site checks. OK.
2017-12-19 | Changed out BAM filter tape beginning at 10:30 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
Leak check as found = 0.4, as left = 0.5.
Conducted site checks. OK.
2017-12-22 | Conducted site checks. OK.
2017-12-26 | Conducted site checks. OK.
Temperature inside the shelter seem a little cool. Turned up temperature on heater.
2017-12-29 | Conducted site checks. OK.
2018-01-03 | Performed QC checks beginning at 11:30 CST.
Conducted site checks. OK.
2018-01-08 | Power failure from 0:00 to 09:00 CST due to a blown fuse.
2018-01-09 | Conducted site checks. OK.
2018-01-10 | Power failure from 2:00 to 14:00 CST due to a blown fuse.
Performed an inspection of the BAM following fuse replacement.
2018-01-12 | Conducted site checks. OK.
2018-01-17 | Conducted site checks. OK.
2018-01-19 | Performed QC checks beginning at 10:10 CST.
Conducted site checks. OK.
2018-01-22 | Conducted site checks. OK.
2018-01-25 | Conducted site checks. OK.
2018-01-29 | Performed QC checks beginning at 10:25 CST.
Conducted site checks. OK.
2018-02-02 | Conducted site checks. OK.
Auxiliary heater stopped working. Will replace on next visit.
2018-02-06 | Conducted site checks. OK.
Replaced auxiliary heater.
2018-02-13 | Conducted site checks. OK.
2018-02-16 | Conducted site checks. OK.
2018-02-20 | Performed QC checks beginning at 11:10 CST.
Conducted site checks. OK.
2018-02-23 | Changed out BAM filter tape beginning at 09:10 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
2018-02-26 | Conducted site checks. OK.
2018-03-01 | Conducted site checks. OK.
2018-03-06 | Performed QC checks beginning at 10:15 CST.
Conducted site checks. OK.
2018-03-09 | Conducted site checks. OK.
2018-03-13 | Conducted site checks. OK.
2018-03-16 | Conducted site checks. OK.
2018-03-20 | Performed QC checks beginning at 09:40 CST.
Conducted site checks. OK.
2018-03-23 | Conducted site checks. OK.
2018-03-27 | Conducted site checks. OK.
Took BAM offline at 09:45 CST to adjust the flows.
As found: 15.0=15.25, 18:4=19.21, and 16.7=17.13.
As left: 15.0=14.82, 18.4=18.35, and 16.7=16.68.
2018-03-29 | J. Sorkness performed a quarterly audit beginning at 08:50 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2018-04-03 | Conducted site checks. OK.
2018-04-06 | Conducted site checks. OK.
2018-04-10 | Performed QC checks beginning at 09:25 CST.
Conducted site checks. OK.
2018-04-13 | Conducted site checks. OK.
2018-04-16 | Conducted site checks. OK.
2018-04-19 | Conducted site checks. OK.
2018-04-23 | Performed QC checks beginning at 09:35 CST.
Conducted site checks. OK.
2018-04-26 | Changed out BAM filter tape beginning at 09:35 CST.
Leak check as found = 0.4, as left = 0.5.
Conducted site checks. OK.
2018-05-01 | Conducted site checks. OK.
2018-05-04 | Conducted site checks. OK.
2018-05-11 | Performed QC checks beginning at 12:20 CST.
Conducted site checks. OK.
2018-05-18 | Conducted site checks. OK.
2018-05-22 | Conducted site checks. OK.
Cleaned enclosure AC filters.
2018-05-25 | Performed QC checks beginning at 09:10 CST.
Conducted site checks. OK.
2018-05-29 | Conducted site checks. OK.
2018-06-01 | Conducted site checks. OK.
2018-06-05 | Conducted site checks. OK.
Cleaned enclosure AC filter.
2018-06-08 | Conducted site checks. OK.
2018-06-11 | Performed QC checks beginning at 14:15 CST.
Conducted site checks. OK.
The monitor recorded flow errors from 11:00 to 12:00 CST likely due to heavy rains.
2018-06-14 | Conducted site checks. OK.
2018-06-18 | Conducted site checks. OK.
2018-06-21 | Conducted site checks. OK.
The monitor recorded a flow error for hour 19:00 CST likely due to heavy rains.
2018-06-22 | The monitor recorded flow errors for hours 08:00 and 13:00 CST likely due to heavy rains.
2018-06-23 | The monitor recorded a flow error for hour 01:00 likely due to heavy rains.
2018-06-25 | Performed QC checks beginning at 09:35 CST.
Conducted site checks. OK.
2018-06-28 | J. Sorkness performed a quarterly audit beginning at 09:10 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
Cleaned enclosure AC filter.
2018-07-02 | Changed out BAM filter tape beginning at 09:10 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As left leak check = 0.3.
2018-07-05 | Cleaned BAM PM10 head.
Conducted site checks. OK.
2018-07-10 | Performed QC checks beginning at 09:45 CST.
Conducted site checks. OK.
2018-07-13 | Conducted site checks. OK.
2018-07-16 | Conducted site checks. OK.
Changed time from 09:52:10 to 09:51 CST.
2018-07-19 | Conducted site checks. OK.
2018-07-24 | Conducted site checks. OK.
2018-07-27 | Performed QC checks beginning at 09:25 CST.
Conducted site checks. OK.
2018-07-30 | Conducted site checks. OK.
2018-08-03 | Conducted site checks. OK.
2018-08-07 | Conducted site checks. OK.
2018-08-10 | Performed QC checks beginning at 09:10 CST.
Conducted site checks. OK.
2018-08-14 | Conducted site checks. OK.
Cleaned enclosure AC filter.
2018-08-15 | Power failure at 16:00 CST.
2018-08-17 | Conducted site checks. OK.
2018-08-23 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2018-08-28 | Conducted site checks. OK.
2018-08-30 | Conducted site checks. OK.
2018-09-04 | Conducted site checks. OK.
2018-09-07 | Changed out BAM filter tape beginning at 10:40 CST.
Leak check as found = 0.3, as left = 0.3.
Conducted site checks. OK.
2018-09-11 | Conducted site checks. OK.
2018-09-14 | Performed QC checks beginning at 10:35 CST.
Conducted site checks. OK.
2018-09-18 | Conducted site checks. OK.
2018-09-21 | Cleaned PM10 sampler head.
Conducted site checks. OK.
2018-09-24 | Performed QC checks beginning at 10:15 CST.
Conducted site checks. OK.
2018-09-27 | J. Sorkness performed a quarterly audit beginning at 10:10 CST, except for PM10 flow.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2018-09-28 | Performed PM10 audit flow checks with DeltaCal s/n 1301 beginning at 08:45 CST.
Results indicate the BAM flows to be within project limits.
2018-10-02 | Conducted site checks. OK.
Changed time from 10:39:05 to 10:38 CST.
2018-10-05 | Conducted site checks. OK.
2018-10-12 | Performed QC checks beginning at 08:20 CST.
Conducted site checks. OK.
2018-10-16 | Conducted site checks. OK.
2018-10-19 | Conducted site checks. OK.
2018-10-23 | Conducted site checks. OK.
2018-10-26 | Performed QC checks beginning at 09:45 CST.
Conducted site checks. OK.
2018-10-30 | Conducted site checks. OK.
Cleaned enclosure AC filter.
2018-11-02 | Conducted site checks. OK.
2018-11-06 | Conducted site checks. OK.
2018-11-12 | Changed out BAM filter tape beginning at 10:10 CST.
Leak check as found = 0.7, as left = 0.5.
Conducted site checks. OK.
2018-11-16 | Performed QC checks beginning at 10:05 CST.
Conducted site checks. OK.
2018-11-27 | Conducted site checks. OK.
2018-11-30 | Performed QC checks beginning at 10:30 CST.
Conducted site checks. OK.
2018-12-03 | Conducted site checks. OK.
2018-12-04 | Tape error message at 11:00 invalidating the hour. Will check on this tomorrow.
2018-12-05 | Onsite to investigate a tape error message on 12/4 at 11:00.
Found the tape to be in good condition and the BAM to be operating normally.
2018-12-06 | J. Sorkness performed a quarterly audit beginning at 10:20 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2018-12-11 | Changed time from 11:13:05 to 11:12 CST.
Conducted site checks. OK.
2018-12-14 | Performed QC checks beginning at 11:40 CST.
Conducted site checks. OK.
2018-12-18 | Conducted site checks. OK.
2018-12-21 | Conducted site checks. OK.
Cleaned PM10 head.
2018-12-28 | Performed QC checks beginning at 10:50 CST.
Conducted site checks. OK.
2018-12-31 | Conducted site checks. OK.
2019-01-08 | Conducted site checks. OK.
2019-01-11 | Performed QC checks beginning at 11:25 CST.
Conducted site checks. OK.
2019-01-15 | Conducted site checks. OK.
2019-01-18 | Changed out BAM filter tape beginning at 10:25 CST.
Leak check as found = 0.4, as left = 0.5.
Conducted site checks. OK.
2019-01-24 | Conducted site checks. OK.
2019-02-01 | Conducted site checks. OK.
2019-02-07 | Performed QC checks beginning at 13:15 CST.
Conducted site checks. OK.
2019-02-11 | Conducted site checks. OK.
2019-02-12 | Hours 06:00 and 09:00 were invalid due to instrument malfunction.
2019-02-15 | Conducted site checks. OK.
2019-02-21 | Performed QC checks beginning at 13:45 CST.
Conducted site checks. OK.
Changed time from 10:34:05 to 10:33 CST.
2019-02-25 | Conducted site checks. OK.
2019-03-01 | Conducted site checks. OK.
2019-03-06 | Performed QC checks beginning at 12:15 CST.
Conducted site checks. OK.
2019-03-12 | Conducted site checks. OK.
2019-03-15 | Trimmed the end of the sample pump line at 09:30 CST.
Conducted site checks. OK.
2019-03-18 | Performed a leak check at 09:50 CST.
As found leak check = 0.4.
Conducted site checks. OK.
2019-03-20 | M. Wolfskill performed a quarterly audit beginning at 08:50 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2019-03-25 | Performed QC checks beginning at 09:30 CST.
Changed out BAM filter tape.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As found leak check = 0.4, As left leak check = 0.4.
2019-03-29 | Conducted site checks. OK.
2019-04-02 | Conducted site checks. OK.
Cleaned PM10 head.
2019-04-05 | Conducted site checks. OK.
2019-04-09 | Conducted site checks. OK.
2019-04-12 | Performed QC checks beginning at 09:20 CST.
Conducted site checks. OK.
2019-04-16 | Conducted site checks. OK.
Inspected enclosure HVAC filter. Found it clean.
2019-04-18 | Conducted site checks. OK.
2019-04-23 | Conducted site checks. OK.
2019-04-26 | Conducted site checks. OK.
2019-04-30 | Performed QC checks beginning at 10:20 CST.
Conducted site checks. OK.
2019-05-03 | Conducted site checks. OK.
2019-05-06 | Conducted site checks. OK.
2019-05-10 | Conducted site checks. OK.
Changed time from 09:27:05 to 09:26 CST.
2019-05-16 | Performed QC checks beginning at 09:20 CST.
Conducted site checks. OK.
2019-05-22 | Conducted site checks. OK.
2019-05-24 | Conducted site checks. OK.
2019-05-28 | Performed QC checks beginning at 10:30 CST.
Changed out BAM filter tape.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As found leak check = 0.3, As left leak check = 0.3.
Conducted site checks. OK.
2019-06-04 | Cleaned PM10 sampler head.
Conducted site checks. OK.
2019-06-11 | Performed QC checks beginning at 10:35 CST.
Conducted site checks. OK.
2019-06-14 | Conducted site checks. OK.
2019-06-18 | Cleaned shelter HVAC filter.
Conducted site checks. OK.
2019-06-21 | Conducted site checks. OK.
2019-06-26 | M. Wolfskill performed a quarterly audit beginning at 09:05 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2019-06-28 | Performed QC checks beginning at 09:40 CST.
Conducted site checks. OK.
2019-07-02 | Conducted site checks. OK.
2019-07-05 | Cleaned shelter HVAC filter.
Conducted site checks. OK.
2019-07-09 | Conducted site checks. OK.
2019-07-12 | Performed QC checks beginning at 09:45 CST.
Conducted site checks. OK.
2019-07-16 | Conducted site checks. OK.
2019-07-19 | Conducted site checks. OK.
2019-07-23 | Cleaned PM10 sampler head.
Conducted site checks. OK.
2019-07-26 | Performed QC checks beginning at 09:20 CST.
Conducted site checks. OK.
2019-07-30 | Changed out BAM filter tape at 09:25 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As found leak check = 0.3, As left leak check = 0.3.
Conducted site checks. OK.
2019-08-02 | Conducted site checks. OK.
2019-08-06 | Conducted site checks. OK.
Changed time from 09:15:05 to 09:14 CST.
2019-08-09 | Performed QC checks beginning at 10:40 CST.
Conducted site checks. OK.
2019-08-13 | Conducted site checks. OK.
2019-08-16 | Conducted site checks. OK.
2019-08-20 | Cleaned shelter HVAC filter.
Conducted site checks. OK.
2019-08-23 | Performed QC checks beginning at 09:20 CST.
Conducted site checks. OK.
2019-08-27 | Conducted site checks. OK.
2019-09-03 | Conducted site checks. OK.
2019-09-05 | Conducted site checks. OK.
2019-09-10 | Performed QC checks beginning at 10:40 CST.
Conducted site checks. OK.
2019-09-11 | Hours 03:00 and 04:00 CST are invalid due to a power failure.
2019-09-13 | Conducted site checks. OK.
2019-09-19 | M. Wolfskill performed a quarterly audit beginning at 07:50 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2019-09-24 | Conducted site checks. OK.
2019-09-27 | Performed QC checks beginning at 09:25 CST.
Conducted site checks. OK.
2019-09-28 | Hour 10:00 CST is invalid due to an instrument malfunction - delta pressure issue.
2019-09-30 | Conducted site checks. OK.
2019-10-04 | Changed out BAM filter tape at 08:45 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As found leak check = 0.3, As left leak check = 0.4.
Conducted site checks. OK.
2019-10-07 | Conducted site checks. OK.
2019-10-09 | Conducted site checks. OK.
2019-10-15 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2019-10-16 | Checked on site communications. Found the router not operational. Power cycled the
router, and normal communications resumed.
2019-10-21 | Conducted site checks. OK.
2019-10-24 | Conducted site checks. OK.
Changed time from 09:38:15 to 09:37 CST.
2019-10-28 | We have been receiving filter tape error status flags since 21:00 on 10/27/2019.
Found one of the tape tension rollers stuck to the left side. Moved the roller back and
forth several times and it seems to be moving freely. Reinstalled the filter tape.
As a result, data from 21:00 CST on 10/27 to 09:00 CST on 10/28 are invalid.
Performed QC checks beginning at 09:50 CST.
Hours 18:00 and 19:00 CST are invalid due to a power failure.
Conducted site checks. OK.
2019-10-29 | Checked on site communications. Found the router not operational. Power cycled the
router, and normal communications resumed. Found water inside the BAM door.
Removed the old caulk around the inlet flange, and recaulked the area.

## ST-N
2014-10-07 | Performed BAM QC checks starting at 12:30 CST.
Cleaned PM10 sample head.
Leak check as found was 0.4, so I did not clean the nozzle.
Conducted site checks. OK.
2014-10-10 | Conducted site checks. OK.
2014-10-13 | Conducted site checks. OK.
2014-10-16 | Conducted site checks. OK.
2014-10-21 | Performed BAM QC checks starting at 10:10 CST.
Conducted site checks. OK.
2014-10-24 | Conducted site checks. OK.
2014-10-28 | Conducted site checks. OK.
2014-10-30 | Conducted site checks. OK.
2014-11-03 | Performed BAM QC checks starting at 13:10 CST.
Conducted site checks. OK.
2014-11-05 | Conducted site checks. OK.
2014-11-10 | Conducted site checks. OK.
2014-11-14 | Conducted site checks. OK.
2014-11-17 | Conducted site checks. OK.
2014-11-21 | Performed BAM QC checks starting at 11:05 CST.
Replaced filter tape.
Conducted site checks. OK.
2014-11-26 | Conducted site checks. OK.
2014-12-01 | Conducted site checks. OK.
2014-12-05 | Performed BAM QC checks starting at 10:25 CST.
Conducted site checks. OK.
Cleaned PM10 sample head.
2014-12-08 | J. Sorkness conducted audits of the BAM.
Conducted site checks. OK.
2014-12-12 | Conducted site checks. OK.
2014-12-16 | Conducted site checks. OK.
2014-12-18 | Performed BAM QC checks starting at 10:25 CST.
Conducted site checks. OK.
2014-12-23 | Conducted site checks. OK.
2014-12-30 | Conducted site checks. OK.
2015-01-02 | Conducted site checks. OK.
2015-01-05 | Conducted site checks. OK.
2015-01-08 | Conducted site checks. OK.
2015-01-12 | Perform QC checks beginning at 11:30 CST.
Conducted site checks. OK.
2015-01-13 | Begin annual zeroing at 09:50 CST.
Conducted site checks. OK.
Changed units' background concentration to 0.0000.
2015-01-16 | Conducted site checks. OK.
Concluded annual zeroing at 13:15 CST.
Changed units' background concentration to -0.0048.
2015-01-20 | Input new zero background concentration of -0.0054 into BAM at 10:40 CST.
Conducted site checks. OK.
2015-01-23 | Conducted site checks. OK.
2015-01-26 | Perform QC checks beginning at 10:45 CST.
Changed time from 10:47:45 to 10:49 CST.
Conducted site checks. OK.
2015-01-29 | Conducted site checks. OK.
2015-02-04 | Conducted site checks. OK.
2015-02-06 | Conducted site checks. OK.
2015-02-10 | Perform QC checks beginning at 12:25 CST.
Conducted site checks. OK.
2015-02-13 | Conducted site checks. OK.
2015-02-16 | Conducted site checks. OK.
2015-02-19 | Conducted site checks. OK.
2015-02-24 | Conducted site checks. OK.
2015-02-27 | Perform QC checks beginning at 12:25 CST.
Conducted site checks. OK.
2015-03-02 | Conducted site checks. OK.
2015-03-06 | Changed filter tape beginning at 13:10 CST.
Performed leak check = 0.5.
Cleaned PM10 head.
Conducted site checks. OK.
2015-03-09 | Conducted site checks. OK.
2015-03-11 | Conducted site checks. OK.
2015-03-17 | Conducted site checks. OK.
2015-03-19 | J. Sorkness performed audits beginning at 12:10 CST.
2015-03-20 | Conducted site checks. OK.
2015-03-24 | Perform QC checks beginning at 10:40 CST.
Changed time from 10:38:55 to 10:40 CST
Conducted site checks. OK.
2015-03-27 | Conducted site checks. OK.
2015-03-31 | Perform leak checks beginning at 11:10 CST.
Leak as found = 0.5, as left = 0.5.
Conducted site checks. OK.
2015-04-01 | Conducted site checks. OK.
2015-04-08 | Perform QC checks beginning at 11:45 CST.
Conducted site checks. OK.
2015-04-10 | Conducted site checks. OK.
2015-04-15 | Conducted site checks. OK.
Turned off power to the BAM at 11:00 CST to install new status reporting wiring.
2015-04-17 | Conducted site checks. OK.
Tested the new status software and wiring. Created a power failure at 11:00 CST.
Created a maintenance status at 12:00 CST. Ben has indicated the initial tests were
positive.
2015-04-24 | Perform QC checks beginning at 10:35 CST.
Conducted site checks. OK.
Tested new status system for flow and tape break errors during hour 13:00 CST.
2015-04-27 | Conducted site checks. OK.
Installed new alarm wiring beginning at 11:50 CST.
2015-04-30 | Conducted site checks. OK.
Changed filter tape beginning at 11:40 CST.
2015-05-04 | Perform QC checks beginning at 10:25 CST.
Conducted site checks. OK.
2015-05-07 | Conducted site checks. OK.
2015-05-11 | Conducted site checks. OK.
2015-05-14 | Conducted site checks. OK.
2015-05-18 | Conducted site checks. OK.
2015-05-21 | Perform QC checks beginning at 11:10 CST.
Conducted site checks. OK.
2015-05-27 | Conducted site checks. OK.
2015-05-29 | Conducted site checks. OK.
2015-06-01 | Conducted site checks. OK.
2015-06-04 | Perform QC checks beginning at 11:45 CST.
Cleaned nozzle and screen.
Cleaned PM10 head.
Conducted site checks. OK.
2015-06-09 | Conducted site checks. OK.
2015-06-12 | Conducted site checks. OK.
2015-06-15 | Performed leak check beginning at 12:10 CST.
Leak check = 0.4.
Conducted site checks. OK.
2015-06-17 | J. Sorkness performed audit checks beginning at 07:40 CST.
2015-06-18 | Conducted site checks. OK.
2015-06-22 | Conducted site checks. OK.
2015-06-24 | Conducted site checks. OK.
Changed filter tape at 07:40 CST.
2015-06-30 | Conducted site checks. OK.
Changed time from 09:46:10 to 09:44 CST.
2015-07-02 | Performed QC checks beginning at 11:25 CST.
Conducted site checks. OK.
2015-07-07 | Conducted site checks. OK.
2015-07-10 | Conducted site checks. OK.
2015-07-14 | Performed QC checks beginning at 09:15 CST.
Conducted site checks. OK.
2015-07-16 | Conducted site checks. OK.
2015-07-22 | Conducted site checks. OK.
2015-07-24 | Conducted site checks. OK.
2015-07-28 | Performed QC checks beginning at 10:30 CST.
Conducted site checks. OK.
2015-07-30 | Conducted site checks. OK.
2015-08-05 | Conducted site checks. OK.
2015-08-07 | Conducted site checks. OK.
2015-08-11 | Performed a BAM leak check beginning at 11:10 CST.
Conducted site checks. OK.
2015-08-12 | J. Sorkness performed a quarterly audit beginning at 07:25 CST.
Conducted site checks. OK.
2015-08-14 | Conducted site checks. OK.
2015-08-18 | Power failure caused by a lightning strike at 19:00 CST.
2015-08-19 | Power restored. Valid data collection resumed at 18:00 CST.
2015-08-21 | Performed QC checks beginning at 11:15 CST.
Changed BAM filter tape.
Conducted site checks. OK.
2015-08-28 | Conducted site checks. OK.
Took BAM offline to record Last M and ABS readings at 10:10 CST.
Last M = 0.824, ABS = 0.806
2015-09-01 | Conducted site checks. OK.
2015-09-04 | Performed QC checks beginning at 08:50 CST.
Conducted site checks. OK.
2015-09-09 | Conducted site checks. OK.
BAM was found in default operate mode, indicating a power failure.
2015-09-11 | Conducted site checks. OK.
Cleaned PM10 head.
2015-09-15 | Conducted site checks. OK.
2015-09-16 | Replaced the sample pump with a rebuilt one. Hours 09:00 and 10:00 CST are invalid due
to the change out.
Conducted site checks. OK.
2015-09-22 | Performed QC checks beginning at 09:40 CST.
Began 72-hour zero calibration following QC checks.
Changed zero offset to 0.0000.
Conducted site checks. OK.
2015-09-25 | Concluded zero calibration.
Adjusted zero offset to former number of -0.0054.
New zero offset numbers will be input next week.
Valid data collection resumed at 11:00 CST.
Conducted site checks. OK.
2015-09-28 | Placed new zero offset into monitor at 10:45 CST.
The new zero offset is -0.0040.
Conducted site checks. OK.
2015-09-30 | Conducted site checks. OK.
2015-10-02 | Conducted site checks. OK.
2015-10-05 | Performed QC checks beginning at 12:05 CST.
Conducted site checks. OK.
2015-10-06 | KCBX power failure from 02:00 to 03:00 CST.
2015-10-08 | Conducted site checks. OK.
2015-10-13 | Conducted site checks. OK.
2015-10-15 | Conducted site checks. OK.
2015-10-19 | Conducted site checks. OK.
2015-10-22 | Performed QC checks beginning at 12:25 CST.
Changed filter tape.
Conducted site checks. OK.
2015-10-27 | Conducted site checks. OK.
2015-10-29 | Conducted site checks. OK.
2015-11-02 | Conducted site checks. OK.
Changed time from 12:40:50 to 12:42 CST.
2015-11-04 | Performed QC checks beginning at 14:45 CST.
Conducted site checks. OK.
2015-11-10 | Conducted site checks. OK.
2015-11-12 | Conducted site checks. OK.
2015-11-16 | Performed a leak check beginning at 12:40 CST.
Leak = 0.5.
Conducted site checks. OK.
2015-11-18 | J. Sorkness performed a quarterly audit beginning at 09:10 CST.
Conducted site checks. OK.
2015-11-23 | Conducted site checks. OK.
2015-11-25 | Conducted site checks. OK.
2015-11-30 | Performed QC checks beginning at 12:50 CST.
Conducted site checks. OK.
2015-12-03 | Conducted site checks. OK.
2015-12-08 | Conducted site checks. OK.
2015-12-10 | Conducted site checks. OK.
Cleaned PM10 head.
2015-12-14 | Conducted site checks. OK.
2015-12-16 | Performed QC checks beginning at 11:50 CST.
Conducted site checks. OK.
2015-12-22 | Changed filter tape beginning at 11:30 CST.
Changed time from 11:26:00 to 11:27 CST.
Conducted site checks. OK.
2015-12-29 | Performed QC checks beginning at 10:40 CST.
Conducted site checks. OK.
2015-12-31 | Conducted site checks. OK.
2016-01-04 | Conducted site checks. OK.
2016-01-06 | Conducted site checks. OK.
2016-01-11 | Conducted site checks. OK.
2016-01-14 | Conducted site checks. OK.
2016-01-19 | Performed QC checks beginning at 11:05 CST.
Conducted site checks. OK.
2016-01-21 | Conducted site checks. OK.
2016-01-25 | Conducted site checks. OK.
2016-01-27 | Conducted site checks. OK.
2016-02-01 | Conducted site checks. OK.
2016-02-04 | Conducted site checks. OK.
2016-02-08 | Performed QC checks beginning at 09:50 CST.
Adjusted temperature from -1.1 to 0.0 degrees C.
Changed filter tape.
Conducted site checks. OK.
2016-02-11 | Conducted site checks. OK.
2016-02-16 | Conducted site checks. OK.
2016-02-19 | Conducted site checks. OK.
2016-02-23 | Performed QC checks beginning at 10:15 CST.
Conducted site checks. OK.
2016-02-26 | Conducted site checks. OK.
2016-02-29 | Conducted site checks. OK.
2016-03-04 | Conducted site checks. OK.
Changed time from 09:44:50 to 09:46 CST.
2016-03-08 | Performed QC checks beginning at 10:10 CST.
Conducted site checks. OK.
2016-03-11 | Conducted site checks. OK.
Cleaned PM10 head.
2016-03-14 | Conducted site checks. OK.
2016-03-17 | Conducted site checks. OK.
Cleaned enclosure air conditioner and changed filter.
2016-03-23 | J. Sorkness performed a quarterly audit beginning at 10:20 CST.
Conducted site checks. OK.
2016-03-25 | Conducted site checks. OK.
2016-04-01 | Conducted site checks. OK.
Changed time from 09:15:00 to 09:14 CST.
2016-04-02 | Power failure caused by high winds starting at 11:00 CST.
2016-04-03 | Power restored to the site at 18:00 CST.
2016-04-04 | Performed QC checks beginning at 09:40 CST.
Conducted site checks. OK.
2016-04-07 | Conducted site checks. OK.
2016-04-11 | Performed preventative maintenace beginning at 10:20 CST.
Cleaned nozzle, screen, and tape rollers.
Changed filter tape.
Conducted site checks. OK.
2016-04-14 | Conducted site checks. OK.
2016-04-18 | Conducted site checks. OK.
2016-04-21 | Performed QC checks beginning at 10:10 CST.
Conducted site checks. OK.
2016-04-25 | Conducted site checks. OK.
2016-04-29 | Conducted site checks. OK.
Cleaned PM10 head.
2016-05-03 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
Last M = 0.832, ABS = 0.808.
2016-05-05 | Conducted site checks. OK.
2016-05-09 | Conducted site checks. OK.
2016-05-12 | Conducted site checks. OK.
Changed time from 11:11:55 to 11:13 CST.
2016-05-16 | Performed QC checks beginning at 10:10 CST.
Conducted site checks. OK.
2016-05-19 | Conducted site checks. OK.
2016-05-24 | Conducted site checks. OK.
2016-05-25 | Elevated concentrations recorded due to high winds.
2016-05-27 | Performed QC checks beginning at 08:45 CST.
Conducted site checks. OK.
2016-05-31 | Conducted site checks. OK.
2016-06-02 | Conducted site checks. OK.
Cleaned BAM enclosure air conditioner filter.
2016-06-07 | J. Sorkness performed a quarterly audit beginning at 07:45 CST.
Conducted site checks. OK.
2016-06-10 | Conducted site checks. OK.
Changed out BAM filter tape at 06:40 CST,
2016-06-14 | Performed QC checks beginning at 13:20 CST.
Conducted site checks. OK.
2016-06-20 | Cleaned PM10 head.
Conducted site checks. OK.
2016-06-23 | Cleaned enclosure air conditioning filter.
Conducted site checks. OK.
2016-06-27 | Performed QC checks beginning at 08:25 CST.
Conducted site checks. OK.
2016-06-30 | Conducted site checks. OK.
2016-07-06 | Conducted site checks. OK.
2016-07-08 | Cleaned PM10 head.
Conducted site checks. OK.
Changed time from 08:55:55 to 08:57 CST.
2016-07-11 | Performed QC checks beginning at 10:20 CST.
Conducted site checks. OK.
2016-07-14 | Conducted site checks. OK.
2016-07-19 | Conducted site checks. OK.
2016-07-22 | Conducted site checks. OK.
2016-07-26 | Performed QC checks beginning at 09:10 CST.
Conducted site checks. OK.
2016-07-28 | Changed out filter tape beginning at 11:30 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
Conducted site checks. OK.
2016-08-02 | Conducted site checks. OK.
2016-08-05 | Conducted site checks. OK.
2016-08-09 | Performed QC checks beginning at 10:10 CST.
Conducted site checks. OK.
2016-08-12 | Conducted site checks. OK.
2016-08-15 | Conducted site checks. OK.
2016-08-18 | Conducted site checks. OK.
2016-08-22 | Performed QC checks beginning at 09:45 CST.
Conducted site checks. OK.
2016-08-26 | Conducted site checks. OK.
Cleaned shelter air condtioner filter.
2016-08-29 | Conducted site checks. OK.
Changed time from 09:25:55 to 09:27 CST.
2016-09-01 | J. Sorkness performed a quarterly audit beginning at 07:55 CST.
Conducted site checks. OK.
2016-09-06 | Conducted site checks. OK.
2016-09-09 | Conducted site checks. OK.
Cleaned PM10 head.
2016-09-13 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2016-09-16 | Conducted site checks. OK.
2016-09-19 | Conducted site checks. OK.
2016-09-22 | Conducted site checks. OK.
Changed out filter tape starting at 09:20 CST.
2016-10-01 | Conducted site checks. OK.
2016-10-04 | Conducted site checks. OK.
2016-10-07 | Conducted site checks. OK.
2016-10-11 | Performed QC checks beginning at 10:40 CST.
Conducted site checks. OK.
2016-10-14 | Conducted site checks. OK.
Cleaned PM10 sample head.
2016-10-17 | Conducted site checks. OK.
2016-10-20 | Conducted site checks. OK.
2016-10-25 | Performed QC checks beginning at 09:25 CST.
Conducted site checks. OK.
Last M = 0.836, ABS = 0.806.
2016-10-28 | Conducted site checks. OK.
Changed time from 09:20:55 to 09:22 CST.
2016-11-02 | Conducted site checks. OK.
2016-11-04 | Conducted site checks. OK.
Cleaned enclosure AC filter.
2016-11-08 | Conducted site checks. OK.
2016-11-11 | QC checks were scheduled, but could not be performed because the BAM was
reading a concentration >100 ug/m3.
Conducted site checks. OK.
2016-11-15 | Performed QC checks beginning at 09:20 CST.
Conducted site checks. OK.
2016-11-18 | Conducted site checks. OK.
2016-11-21 | Hours 1700 and 1800 CST were invalid due to heavy loading at the filter tape and
subsequent pressure drop, during a fire in the building to the north of the sampler,
as reported by KCBX staff.
2016-11-22 | Performed QC checks beginning at 10:35 CST.
Changed out filter tape.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
Conducted site checks. OK.
2016-11-25 | Conducted site checks. OK.
2016-11-30 | J. Sorkness performed a quarterly audit beginning at 15:45 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2016-12-02 | Conducted site checks. OK.
2016-12-06 | Conducted site checks. OK.
2016-12-09 | Conducted site checks. OK.
2016-12-13 | Performed QC checks beginning at 10:25 CST.
Conducted site checks. OK.
2016-12-16 | Conducted site checks. OK.
2016-12-20 | Conducted site checks. OK.
2016-12-22 | Conducted site checks. OK.
Cleaned PM10 head.
2016-12-28 | Performed QC checks beginning at 12:45 CST.
Conducted site checks. OK.
2016-12-30 | Conducted site checks. OK.
Changed time from 10:17:55 to 10:19 CST.
2017-01-03 | Conducted site checks. OK.
2017-01-06 | Conducted site checks. OK.
2017-01-09 | Conducted site checks. OK.
2017-01-13 | Performed QC checks beginning at 10:30 CST.
Conducted site checks. OK.
2017-01-17 | Performed filter tape change out beginning at 10:45 CST.
Conducted site checks. OK.
2017-01-20 | Power failure at 8:00 CST.
2017-01-23 | Conducted site checks. OK.
2017-01-27 | Performed QC checks beginning at 10:25 CST.
Conducted site checks. OK.
2017-01-31 | Conducted site checks. OK.
2017-02-06 | Conducted site checks. OK.
2017-02-13 | Performed QC checks beginning at 10:35 CST.
Conducted site checks. OK.
2017-02-17 | Conducted site checks. OK.
2017-02-23 | Conducted site checks. OK.
2017-02-27 | Performed QC checks beginning at 10:10 CST.
Conducted site checks. OK.
2017-03-03 | Conducted site checks. OK.
2017-03-08 | Cleaned PM10 heads.
Conducted site checks. OK.
2017-03-16 | Performed QC checks beginning at 09:10 CST.
Conducted site checks. OK.
2017-03-20 | Cleaned enclosure AC filters.
Conducted site checks. OK.
2017-03-24 | J. Sorkness performed a quarterly audit beginning at 09:50 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2017-03-27 | Changed out BAM filter tape beginning at 09:20 CST.
Conducted site checks. OK.
2017-03-31 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2017-04-07 | Conducted site checks. OK.
2017-04-10 | Conducted site checks. OK.
2017-04-13 | Performed QC checks beginning at 09:20 CST.
Conducted site checks. OK.
2017-04-18 | Conducted site checks. OK.
2017-04-21 | Conducted site checks. OK.
2017-04-25 | Performed QC checks beginning at 09:25 CST.
Conducted site checks. OK.
2017-04-28 | Conducted site checks. OK.
2017-05-05 | Conducted site checks. OK.
Changed time from 09:13:55 to 09:15 CST.
2017-05-12 | Performed QC checks beginning at 09:10 CST.
Conducted site checks. OK.
2017-05-14 | Hour 00 has been invalidated by the BAM. D. Szatkowski checked and found a "D" flag
for that hour, which indicates a dirty cal membrane, or a piece of filter paper
that fell on the membrane, which was later dislodged. It appears his diagnosis was correct.
I found a small piece of filter paper on one of the pinch rollers during cleaning on 5/30/17.
2017-05-16 | Conducted site checks. OK.
2017-05-19 | Conducted site checks. OK.
Cleaned PM10 heads.
2017-05-23 | Conducted site checks. OK.
2017-05-26 | Performed QC checks beginning at 08:35 CST.
Conducted site checks. OK.
Last M = 0.842, ABS = 0.806.
2017-05-30 | Changed out filter tape beginning at 09:40 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
Leak check as found = 0.5. Leak check as left = 0.4
Conducted site checks. OK.
2017-06-05 | Conducted site checks. OK.
Hour 00 has been invalidated by the BAM. D. Szatkowski checked and found a "D" flag
for that hour, which indicates a dirty cal membrane. Will pick up a can of compressed air
to clean the membrane.
2017-06-08 | Performed QC checks beginning at 10:20 CST.
Conducted site checks. OK.
2017-06-13 | Conducted site checks. OK.
2017-06-15 | Accidentally power cycled the BAM at 09:00 CST. Power to the BAM was lost while the
power company was replacing the fuse to the ST-SE BAM and meteorological station at
11:00 CST. The KCBX electrician advised me of this ahead of time.
2017-06-20 | Conducted site checks. OK.
2017-06-21 | Hour 00 has been invalidated by the BAM. D. Szatkowski checked and found a "D" flag
for that hour, which indicates a dirty cal membrane.
2017-06-23 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
Cleaned shelter AC filter.
Cleaned block and membrane with compressed air. Started Span Membrane Test at
10:40 CST. Span reading went down from 0.843 to 0.820. ABS=.806.
2017-06-29 | J. Sorkness performed a quarterly audit beginning at 08:40 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2017-07-05 | Conducted site checks. OK.
2017-07-11 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2017-07-14 | Conducted site checks. OK.
2017-07-18 | Conducted site checks. OK.
Cleaned PM10 head.
2017-07-21 | Conducted site checks. OK.
2017-07-24 | Performed QC checks beginning at 09:20 CST.
Changed time from 09:16:50 to 09:18 CST.
Conducted site checks. OK.
2017-07-27 | Conducted site checks. OK.
2017-07-31 | Conducted site checks. OK.
2017-08-04 | Conducted site checks. OK.
Changed out BAM filter tape at 12:10 CST.
2017-08-08 | Performed QC checks beginning at 09:25 CST.
Conducted site checks. OK.
2017-08-11 | Conducted site checks. OK.
2017-08-15 | Conducted site checks. OK.
2017-08-16 | Conducted site checks. OK.
Replaced shelter AC filter.
2017-08-22 | Conducted site checks. OK.
2017-08-25 | Performed QC checks beginning at 09:15 CST.
Conducted site checks. OK.
2017-08-29 | Conducted site checks. OK.
2017-08-31 | Conducted site checks. OK.
2017-09-05 | Conducted site checks. OK.
Cleaned PM10 head.
2017-09-08 | J. Sorkness performed a quarterly audit beginning at 10:10 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2017-09-12 | Performed QC checks beginning at 09:35 CST.
Conducted site checks. OK.
2017-09-15 | Conducted site checks. OK.
2017-09-19 | Conducted site checks. OK.
2017-09-22 | Performed QC checks beginning at 09:20 CST.
Conducted site checks. OK.
2017-10-02 | Conducted site checks. OK.
2017-10-05 | Performed QC checks beginning at 10:25 CST.
Conducted site checks. OK.
2017-10-09 | Changed out BAM filter tape at 10:25.
Conducted site checks. OK.
2017-10-12 | Conducted site checks. OK.
2017-10-16 | Conducted site checks. OK.
2017-10-20 | Conducted site checks. OK.
2017-10-24 | Conducted site checks. OK.
2017-10-27 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2017-11-06 | Conducted site checks. OK.
2017-11-10 | Conducted site checks. OK.
2017-11-14 | Rebuilt the BAM sample pump. Offline at 11:30 CST for the change out.
Conducted site checks. OK.
2017-11-17 | Performed QC checks beginning at 08:25 CST.
Conducted site checks. OK.
2017-11-24 | Conducted site checks. OK.
2017-11-27 | Performed QC checks beginning at 10:40 CST.
Conducted site checks. OK.
2017-11-30 | J. Sorkness performed a quarterly audit beginning at 08:20 CST.
Results indicate all systems to be operating within allowable project limits.
Adjusted temperature from 5.7 to 8.0 degrees C.
Conducted site checks. OK.
2017-12-05 | Conducted site checks. OK.
2017-12-12 | Conducted site checks. OK.
2017-12-15 | Performed QC checks beginning at 10:30 CST.
Conducted site checks. OK.
2017-12-19 | Changed out BAM filter tape beginning at 11:30 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
Leak check as found = 0.4, as left = 0.5.
Conducted site checks. OK.
2017-12-22 | Conducted site checks. OK.
2017-12-26 | Conducted site checks. OK.
2017-12-29 | Conducted site checks. OK.
2018-01-02 | The monitor recorded flow errors for hours 05:00 and 07:00 CST, likely due to extreme low
temperatures. The BAM resumed normal operation at 08:00 CST.
Concentrations were recorded by the BAM. D. Szatkowski will check with MetOne to
see if these concentrations are valid.
2018-01-03 | Performed QC checks beginning at 11:20 CST.
Conducted site checks. OK.
2018-01-09 | Conducted site checks. OK.
Changed time from 10:33:55 to 10:35 CST.
2018-01-12 | Conducted site checks. OK.
2018-01-17 | Conducted site checks. OK.
2018-01-19 | Performed QC checks beginning at 10:20 CST.
Conducted site checks. OK.
2018-01-22 | Conducted site checks. OK.
2018-01-25 | Conducted site checks. OK.
2018-01-29 | Performed QC checks beginning at 10:35 CST.
Conducted site checks. OK.
2018-02-02 | Conducted site checks. OK.
2018-02-06 | Conducted site checks. OK.
2018-02-13 | Conducted site checks. OK.
2018-02-16 | Conducted site checks. OK.
2018-02-20 | Performed QC checks beginning at 11:20 CST.
Conducted site checks. OK.
2018-02-23 | Changed out BAM filter tape beginning at 09:25 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
2018-02-26 | Conducted site checks. OK.
2018-03-01 | Conducted site checks. OK.
2018-03-06 | Performed QC checks beginning at 10:35 CST.
Conducted site checks. OK.
2018-03-09 | Conducted site checks. OK.
2018-03-13 | Conducted site checks. OK.
2018-03-16 | Conducted site checks. OK.
Changed time from 09:49:50 to 09:51 CST.
2018-03-20 | Performed QC checks beginning at 10:30 CST.
Conducted site checks. OK.
2018-03-23 | Conducted site checks. OK.
2018-03-27 | Conducted site checks. OK.
2018-03-29 | J. Sorkness performed a quarterly audit beginning at 09:10 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2018-04-03 | Conducted site checks. OK.
2018-04-06 | Conducted site checks. OK.
2018-04-10 | Performed QC checks beginning at 09:35 CST.
Conducted site checks. OK.
2018-04-13 | Conducted site checks. OK.
2018-04-16 | Conducted site checks. OK.
2018-04-19 | Conducted site checks. OK.
2018-04-23 | Performed QC checks beginning at 10:25 CST.
Conducted site checks. OK.
2018-04-26 | Changed out BAM filter tape beginning at 09:50 CST.
Leak check as found = 0.4, as left = 0.5.
Conducted site checks. OK.
2018-05-01 | Conducted site checks. OK.
2018-05-04 | Conducted site checks. OK.
2018-05-11 | Performed QC checks beginning at 12:35 CST.
Conducted site checks. OK.
2018-05-18 | Conducted site checks. OK.
2018-05-22 | Conducted site checks. OK.
Cleaned enclosure AC filters.
2018-05-25 | Performed QC checks beginning at 09:20 CST.
Conducted site checks. OK.
2018-05-29 | Conducted site checks. OK.
2018-06-01 | Conducted site checks. OK.
2018-06-05 | Conducted site checks. OK.
Cleaned enclosure AC filter.
2018-06-08 | Conducted site checks. OK.
2018-06-11 | Performed QC checks beginning at 14:35 CST.
Conducted site checks. OK.
2018-06-14 | Conducted site checks. OK.
2018-06-18 | Conducted site checks. OK.
2018-06-21 | Conducted site checks. OK.
2018-06-25 | Performed QC checks beginning at 09:45 CST.
Conducted site checks. OK.
2018-06-28 | J. Sorkness performed a quarterly audit beginning at 09:25 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
Cleaned enclosure AC filter.
2018-07-02 | Changed out BAM filter tape beginning at 09:20 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As left leak check = 0.5.
2018-07-05 | Cleaned BAM PM10 head.
Conducted site checks. OK.
2018-07-10 | Performed QC checks beginning at 09:35 CST.
Changed time from 09:34:55 to 09:36 CST.
Conducted site checks. OK.
2018-07-13 | Conducted site checks. OK.
2018-07-16 | Conducted site checks. OK.
2018-07-19 | Conducted site checks. OK.
Moved sample pump away from the data logger to reduce shelter temperature.
2018-07-24 | Conducted site checks. OK.
2018-07-27 | Performed QC checks beginning at 09:40 CST.
Conducted site checks. OK.
2018-07-30 | Conducted site checks. OK.
2018-08-03 | Conducted site checks. OK.
2018-08-07 | Conducted site checks. OK.
2018-08-10 | Performed QC checks beginning at 09:20 CST.
Conducted site checks. OK.
2018-08-14 | Conducted site checks. OK.
Cleaned enclosure AC filter.
2018-08-17 | Conducted site checks. OK.
2018-08-23 | Performed QC checks beginning at 09:40 CST.
Conducted site checks. OK.
2018-08-28 | Conducted site checks. OK.
2018-08-30 | Conducted site checks. OK.
2018-09-04 | Conducted site checks. OK.
2018-09-07 | Changed out BAM filter tape beginning at 10:20 CST.
Leak check as found = 0.4, as left = 0.5.
Conducted site checks. OK.
Changed time from 10:20:55 to 10:22 CST.
2018-09-11 | Conducted site checks. OK.
2018-09-14 | Performed QC checks beginning at 10:20 CST.
Conducted site checks. OK.
2018-09-18 | Conducted site checks. OK.
2018-09-21 | Cleaned PM10 sampler head.
Conducted site checks. OK.
2018-09-24 | Performed QC checks beginning at 10:25 CST.
Conducted site checks. OK.
2018-09-27 | J. Sorkness performed a quarterly audit beginning at 10:20 CST, except for PM10 flow.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2018-09-28 | Performed PM10 audit flow checks with DeltaCal s/n 1301 beginning at 08:55 CST.
Results indicate the BAM flows to be within project limits.
2018-10-02 | Conducted site checks. OK.
2018-10-05 | Conducted site checks. OK.
2018-10-12 | Performed QC checks beginning at 08:35 CST.
Conducted site checks. OK.
2018-10-16 | Conducted site checks. OK.
2018-10-19 | Conducted site checks. OK.
2018-10-23 | Conducted site checks. OK.
2018-10-26 | Performed QC checks beginning at 08:40 CST.
Conducted site checks. OK.
2018-10-30 | Conducted site checks. OK.
Cleaned enclosure AC filter.
2018-11-02 | Conducted site checks. OK.
2018-11-06 | Conducted site checks. OK.
2018-11-12 | Changed out BAM filter tape beginning at 10:20 CST.
Leak check as found = 0.6, as left = 0.5.
Conducted site checks. OK.
2018-11-16 | Performed QC checks beginning at 10:15 CST.
Conducted site checks. OK.
2018-11-27 | Conducted site checks. OK.
2018-11-30 | Performed QC checks beginning at 10:15 CST.
Conducted site checks. OK.
2018-12-03 | Conducted site checks. OK.
2018-12-06 | J. Sorkness performed a quarterly audit beginning at 10:45 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2018-12-11 | Conducted site checks. OK.
2018-12-14 | Performed QC checks beginning at 11:30 CST.
Conducted site checks. OK.
2018-12-18 | Conducted site checks. OK.
2018-12-21 | Conducted site checks. OK.
Cleaned PM10 head.
2018-12-28 | Performed QC checks beginning at 10:40 CST.
Conducted site checks. OK.
2018-12-31 | Conducted site checks. OK.
2019-01-08 | Conducted site checks. OK.
2019-01-11 | Performed QC checks beginning at 11:40 CST.
Conducted site checks. OK.
2019-01-15 | Conducted site checks. OK.
2019-01-18 | Changed out BAM filter tape beginning at 10:35 CST.
Leak check as found = 0.5, as left = 0.5.
Conducted site checks. OK.
Changed time from 10:44:55 to 10:46 CST.
2019-01-24 | Conducted site checks. OK.
2019-02-01 | Conducted site checks. OK.
2019-02-07 | Performed QC checks beginning at 13:25 CST.
Conducted site checks. OK.
2019-02-11 | Conducted site checks. OK.
2019-02-15 | Conducted site checks. OK.
2019-02-21 | Performed QC checks beginning at 14:10 CST.
Conducted site checks. OK.
2019-02-25 | Conducted site checks. OK.
2019-03-01 | Conducted site checks. OK.
2019-03-06 | Performed QC checks beginning at 12:30 CST.
Conducted site checks. OK.
2019-03-12 | Conducted site checks. OK.
2019-03-15 | Trimmed the end of the sample pump line at 09:30 CST.
Conducted site checks. OK.
2019-03-18 | Performed a leak check at 09:55 CST.
As found leak check = 0.5.
Conducted site checks. OK.
2019-03-20 | M. Wolfskill performed a quarterly audit beginning at 09:10 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2019-03-25 | Performed QC checks beginning at 09:45 CST.
Changed out BAM filter tape beginning at 09:10 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As found leak check = 0.5, As left leak check = 0.5.
2019-03-29 | Conducted site checks. OK.
2019-04-02 | Conducted site checks. OK.
Cleaned PM10 head.
2019-04-05 | Conducted site checks. OK.
2019-04-09 | Conducted site checks. OK.
2019-04-10 | Power failure during hour 12:00 while KCBX electricians were rewiring the circuit.
2019-04-12 | Performed QC checks beginning at 09:35 CST.
Conducted site checks. OK.
2019-04-16 | Conducted site checks. OK.
Inspected enclosure HVAC filter. Found it clean.
2019-04-18 | Conducted site checks. OK.
2019-04-23 | Conducted site checks. OK.
2019-04-26 | Took the BAM offline at 10:00 CST to check the DeltaCal flow.
It was reading 11.5, instead of the expected 15.0.
Conducted site checks. OK.
2019-04-30 | Performed QC checks beginning at 10:35 CST.
Conducted site checks. OK.
2019-05-03 | Conducted site checks. OK.
2019-05-06 | Conducted site checks. OK.
2019-05-10 | Conducted site checks. OK.
2019-05-16 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2019-05-22 | Conducted site checks. OK.
2019-05-24 | Conducted site checks. OK.
2019-05-28 | Performed QC checks beginning at 11:10 CST.
Changed time from 11:06:55 to 11:08 CST.
Changed out BAM filter tape.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As found leak check = 0.4, As left leak check = 0.5.
Conducted site checks. OK.
2019-06-04 | Cleaned PM10 sampler head.
Conducted site checks. OK.
2019-06-11 | Performed QC checks beginning at 10:25 CST.
Conducted site checks. OK.
2019-06-14 | Conducted site checks. OK.
2019-06-18 | Cleaned shelter HVAC filter.
Conducted site checks. OK.
2019-06-21 | Conducted site checks. OK.
2019-06-26 | M. Wolfskill performed a quarterly audit beginning at 09:30 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2019-06-28 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2019-07-02 | Conducted site checks. OK.
2019-07-05 | Cleaned shelter HVAC filter.
Conducted site checks. OK.
2019-07-09 | Conducted site checks. OK.
2019-07-12 | Performed QC checks beginning at 09:35 CST.
Conducted site checks. OK.
2019-07-16 | Conducted site checks. OK.
2019-07-19 | Changed time from 08:37:55 to 08:39 CST.
Conducted site checks. OK.
2019-07-23 | Cleaned PM10 sampler head.
Conducted site checks. OK.
2019-07-26 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2019-07-30 | Changed out BAM filter tape at 09:35 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As found leak check = 0.4, As left leak check = 0.4.
Conducted site checks. OK.
2019-08-02 | Conducted site checks. OK.
2019-08-06 | Conducted site checks. OK.
2019-08-09 | Performed QC checks beginning at 10:25 CST.
Conducted site checks. OK.
2019-08-13 | Conducted site checks. OK.
2019-08-16 | Conducted site checks. OK.
2019-08-20 | Cleaned shelter HVAC filter.
Conducted site checks. OK.
2019-08-23 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2019-08-27 | Conducted site checks. OK.
2019-09-03 | Conducted site checks. OK.
2019-09-05 | Conducted site checks. OK.
2019-09-10 | Performed QC checks beginning at 10:30 CST.
Conducted site checks. OK.
2019-09-11 | Hours 03:00 and 04:00 CST are invalid due to a power failure.
2019-09-13 | Conducted site checks. OK.
Changed time from 09:54:50 to 09:56 CST.
2019-09-19 | M. Wolfskill performed a quarterly audit beginning at 07:40 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2019-09-24 | Conducted site checks. OK.
2019-09-27 | Performed QC checks beginning at 10:30 CST.
Conducted site checks. OK.
2019-09-30 | Conducted site checks. OK.
2019-10-04 | Changed out BAM filter tape at 09:10 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As found leak check = 0.5, As left leak check = 0.5.
Conducted site checks. OK.
2019-10-07 | Conducted site checks. OK.
2019-10-09 | Conducted site checks. OK.
2019-10-15 | Performed QC checks beginning at 09:40 CST.
Conducted site checks. OK.
2019-10-21 | Conducted site checks. OK.
2019-10-24 | Conducted site checks. OK.
2019-10-28 | Performed QC checks beginning at 11:40 CST.
Changed time from 11:38:55 to 11:40 CST.
Conducted site checks. OK.

## ST-SW
2014-10-07 | Performed BAM QC checks starting at 12:50 CST.
Cleaned PM10 sample head.
Leak check as found was 0.3, so I did not clean the nozzle.
Conducted site checks. OK.
2014-10-10 | Conducted site checks. OK.
2014-10-13 | Conducted site checks. OK.
2014-10-16 | Conducted site checks. OK.
2014-10-21 | Performed BAM QC checks starting at 10:25 CST.
Changed time from 10:28:05 to 10:27.
Conducted site checks. OK.
2014-10-24 | Conducted site checks. OK.
2014-10-28 | Conducted site checks. OK.
2014-10-30 | Conducted site checks. OK.
2014-11-03 | Performed BAM QC checks starting at 13:20 CST.
Conducted site checks. OK.
2014-11-05 | Conducted site checks. OK.
2014-11-10 | Conducted site checks. OK.
2014-11-14 | Conducted site checks. OK.
2014-11-17 | Conducted site checks. OK.
2014-11-21 | Performed BAM QC checks starting at 11:25 CST.
Replaced filter tape.
Conducted site checks. OK.
2014-11-26 | Conducted site checks. OK.
2014-12-01 | Conducted site checks. OK.
2014-12-05 | Performed BAM QC checks starting at 11:30 CST.
Changed time from 11:31:05 to 11:30 CST.
Conducted site checks. OK.
Cleaned PM10 sample head.
2014-12-08 | J. Sorkness conducted audits of the BAM.
Conducted site checks. OK.
2014-12-12 | Conducted site checks. OK.
2014-12-16 | Conducted site checks. OK.
2014-12-18 | Performed BAM QC checks starting at 10:40 CST.
Conducted site checks. OK.
2014-12-23 | Conducted site checks. OK.
2014-12-30 | Conducted site checks. OK.
2015-01-02 | Conducted site checks. OK.
2015-01-05 | Conducted site checks. OK.
2015-01-08 | Conducted site checks. OK.
2015-01-12 | Perform QC checks beginning at 11:50 CST.
Conducted site checks. OK.
2015-01-13 | Begin annual zeroing at 10:05 CST.
Conducted site checks. OK.
Changed units' background concentration to 0.0000.
2015-01-16 | Conducted site checks. OK.
Concluded annual zeroing at 11:20 CST.
Changed units' background concentration to -0.0034.
2015-01-20 | Input new zero background concentration of -0.0035 into BAM at 10:50 CST.
Conducted site checks. OK.
2015-01-23 | Conducted site checks. OK.
2015-01-26 | Perform QC checks beginning at 11:00 CST.
Conducted site checks. OK.
2015-01-29 | Conducted site checks. OK.
2015-02-04 | Conducted site checks. OK.
2015-02-06 | Conducted site checks. OK.
2015-02-10 | Perform QC checks beginning at 12:40 CST.
Conducted site checks. OK.
2015-02-12 | Power failure from 07:40 to 08:05 CST. Checked with Todd Chandler of KCBX and he
indicated the PF was caused by an electrician accidentally turning off the circuit breaker
to the shelter during a lock-out tag-out operation, while working on the pumps.
2015-02-13 | Conducted site checks. OK.
2015-02-16 | Conducted site checks. OK.
2015-02-19 | Conducted site checks. OK.
2015-02-24 | Conducted site checks. OK.
2015-02-25 | Power failure from 08:10 to 08:40 due to an electrician lockout/tagout issue.
2015-02-27 | Perform QC checks beginning at 12:40 CST.
Conducted site checks. OK.
2015-03-02 | Conducted site checks. OK.
2015-03-06 | Changed filter tape beginning at 13:40 CST.
Performed leak check = 0.5.
Cleaned PM10 head.
Conducted site checks. OK.
2015-03-09 | Conducted site checks. OK.
2015-03-11 | Conducted site checks. OK.
2015-03-17 | Conducted site checks. OK.
2015-03-19 | J. Sorkness performed audits beginning at 12:20 CST.
2015-03-20 | Conducted site checks. OK.
2015-03-24 | Perform QC checks beginning at 10:50 CST.
Conducted site checks. OK.
2015-03-27 | Conducted site checks. OK.
2015-03-31 | Perform leak checks beginning at 11:30 CST.
Leak as found = 0.4, as left = 0.5.
Conducted site checks. OK.
Changed time from 11:33:05 to 11:32 CST.
Mistakenly took the BAM offline when the previous hour's concentration was 200 ug/m3.
Based on north winds at the time, this concentration appears valid.
2015-04-01 | Conducted site checks. OK.
2015-04-08 | Perform QC checks beginning at 12:15 CST.
Conducted site checks. OK.
2015-04-10 | Conducted site checks. OK.
2015-04-15 | Conducted site checks. OK.
2015-04-17 | Conducted site checks. OK.
2015-04-24 | Perform QC checks beginning at 10:45 CST.
Conducted site checks. OK.
2015-04-27 | Conducted site checks. OK.
Installed new alarm wiring beginning at 12:10 CST.
2015-04-30 | Conducted site checks. OK.
Changed filter tape beginning at 13:15 CST.
2015-05-04 | Perform QC checks beginning at 10:40 CST.
Conducted site checks. OK.
2015-05-07 | Conducted site checks. OK.
2015-05-11 | Conducted site checks. OK.
2015-05-14 | Conducted site checks. OK.
Power failure during hour 11:00 CST.
2015-05-15 | Power failure during hour 22:00 CST.
2015-05-18 | Conducted site checks. OK.
2015-05-21 | Perform QC checks beginning at 11:30 CST.
Changed time from 11:28:08 to 11:27 CST.
Conducted site checks. OK.
2015-05-27 | Conducted site checks. OK.
Found the monitor in default operating mode, indicating a power failure.
2015-05-29 | Conducted site checks. OK.
Found the monitor in default operating mode, indicating a power failure.
2015-06-01 | Conducted site checks. OK.
2015-06-04 | Perform QC checks beginning at 12:15 CST.
Cleaned nozzle and screen.
Cleaned PM10 head.
Conducted site checks. OK.
2015-06-09 | Conducted site checks. OK.
2015-06-12 | Conducted site checks. OK.
2015-06-15 | Performed leak check beginning at 12:25 CST.
Leak check = 0.4.
Conducted site checks. OK.
2015-06-17 | J. Sorkness performed audit checks beginning at 08:20 CST.
2015-06-18 | Conducted site checks. OK.
2015-06-22 | Conducted site checks. OK.
2015-06-24 | Conducted site checks. OK.
Changed filter tape at 07:50 CST.
2015-06-30 | Conducted site checks. OK.
2015-07-02 | Performed QC checks beginning at 08:45 CST.
Conducted site checks. OK.
2015-07-07 | Conducted site checks. OK.
2015-07-08 | Power failure from 14:00 to 15:00 caused by KCBX.
2015-07-10 | Conducted site checks. OK.
2015-07-14 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2015-07-16 | Conducted site checks. OK.
2015-07-22 | Conducted site checks. OK.
2015-07-24 | Conducted site checks. OK.
2015-07-28 | Performed QC checks beginning at 10:45 CST.
Conducted site checks. OK.
2015-07-30 | Conducted site checks. OK.
2015-08-05 | Conducted site checks. OK.
Changed time from 09:19:05 to 09:18 CST.
2015-08-07 | Conducted site checks. OK.
2015-08-11 | Performed a BAM leak check beginning at 11:20 CST.
Conducted site checks. OK.
2015-08-12 | J. Sorkness performed a quarterly audit beginning at 08:40 CST.
Conducted site checks. OK.
2015-08-14 | Conducted site checks. OK.
2015-08-18 | Power failure caused by a lightning strike at 19:00 CST.
2015-08-19 | Power restored. Valid data collection resumed at 17:00 CST.
2015-08-21 | Performed QC checks beginning at 11:15 CST.
Changed BAM filter tape.
Conducted site checks. OK.
2015-08-28 | Conducted site checks. OK.
Took BAM offline to record Last M and ABS readings at 10:20 CST.
Last M = 0.848, ABS = 0.839
2015-09-01 | Conducted site checks. OK.
2015-09-04 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2015-09-09 | Conducted site checks. OK.
2015-09-11 | Conducted site checks. OK.
Cleaned PM10 head.
2015-09-15 | Conducted site checks. OK.
2015-09-16 | Replaced the sample pump with a rebuilt one. Hour 10:00 CST is invalid due to the change
out.
Conducted site checks. OK.
2015-09-17 | Hour 06:00 CST is invalid due to a power failure.
2015-09-22 | Performed QC checks beginning at 10:10 CST.
Began 72-hour zero calibration following QC checks.
Changed zero offset to 0.0000.
Conducted site checks. OK.
2015-09-25 | Concluded zero calibration.
Adjusted zero offset to former number of -0.0035.
New zero offset numbers will be input next week.
Valid data collection resumed at 11:00 CST.
Conducted site checks. OK.
2015-09-28 | Placed new zero offset into monitor at 10:55 CST.
The new zero offset is -0.0013.
Conducted site checks. OK.
2015-09-30 | Conducted site checks. OK.
2015-10-02 | Conducted site checks. OK.
2015-10-05 | Performed QC checks beginning at 12:20 CST.
Conducted site checks. OK.
2015-10-08 | Conducted site checks. OK.
Salt trucks on the property to the south are generating visible fugitive dust.
2015-10-13 | Conducted site checks. OK.
2015-10-15 | Conducted site checks. OK.
2015-10-19 | Conducted site checks. OK.
2015-10-22 | Performed QC checks beginning at 13:05 CST.
Changed filter tape.
Conducted site checks. OK.
2015-10-27 | Conducted site checks. OK.
2015-10-29 | Conducted site checks. OK.
2015-11-02 | Conducted site checks. OK.
2015-11-04 | Performed QC checks beginning at 15:05 CST.
Conducted site checks. OK.
2015-11-10 | Conducted site checks. OK.
2015-11-12 | Conducted site checks. OK.
2015-11-16 | A leak check was schedulded, but could not be performed due to a last measured
concentration of 106 ug/m3. A salt barge is being loaded to the south of the monitor.
Conducted site checks. OK.
2015-11-18 | J. Sorkness performed a quarterly audit beginning at 09:25 CST.
Conducted site checks. OK.
2015-11-23 | Conducted site checks. OK.
2015-11-25 | Conducted site checks. OK.
2015-11-30 | Performed QC checks beginning at 13:10 CST.
Conducted site checks. OK.
2015-12-03 | Conducted site checks. OK.
2015-12-08 | Conducted site checks. OK.
2015-12-10 | Conducted site checks. OK.
Cleaned PM10 head.
2015-12-14 | Conducted site checks. OK.
2015-12-16 | Performed QC checks beginning at 12:10 CST.
Conducted site checks. OK.
2015-12-22 | Changed filter tape beginning at 11:40 CST.
Conducted site checks. OK.
2015-12-29 | Performed QC checks beginning at 11:05 CST.
Conducted site checks. OK.
2015-12-31 | Conducted site checks. OK.
2016-01-04 | Conducted site checks. OK.
2016-01-06 | Conducted site checks. OK.
2016-01-11 | Conducted site checks. OK.
2016-01-14 | Conducted site checks. OK.
2016-01-19 | Performed QC checks beginning at 11:20 CST.
Conducted site checks. OK.
2016-01-21 | Conducted site checks. OK.
2016-01-25 | Conducted site checks. OK.
2016-01-27 | Conducted site checks. OK.
2016-02-01 | Conducted site checks. OK.
2016-02-04 | Conducted site checks. OK.
2016-02-08 | Performed QC checks beginning at 10:15 CST.
Changed time from 10:28:10 to 10:27 CST.
Changed filter tape.
Conducted site checks. OK.
2016-02-11 | Conducted site checks. OK.
2016-02-16 | Conducted site checks. OK.
2016-02-19 | Conducted site checks. OK.
2016-02-23 | Performed QC checks beginning at 10:30 CST.
Conducted site checks. OK.
2016-02-26 | Conducted site checks. OK.
2016-02-29 | Conducted site checks. OK.
2016-03-04 | Conducted site checks. OK.
2016-03-08 | Performed QC checks beginning at 10:20 CST.
Conducted site checks. OK.
2016-03-11 | Conducted site checks. OK.
Cleaned PM10 head.
2016-03-14 | Conducted site checks. OK.
2016-03-17 | Conducted site checks. OK.
Cleaned enclosure air conditioner and changed filter.
2016-03-23 | J. Sorkness performed a quarterly audit beginning at 10:35 CST.
Conducted site checks. OK.
2016-03-25 | Conducted site checks. OK.
2016-04-01 | Conducted site checks. OK.
2016-04-02 | Power failure caused by high winds starting at 19:00 CST.
2016-04-03 | Power restored to the site at 10:00 CST.
2016-04-04 | Performed QC checks beginning at 10:15 CST.
Conducted site checks. OK.
2016-04-07 | Conducted site checks. OK.
2016-04-11 | Performed preventative maintenace beginning at 11:10 CST.
Cleaned nozzle, screen, and tape rollers.
Changed filter tape.
Conducted site checks. OK.
2016-04-14 | Conducted site checks. OK.
2016-04-18 | Conducted site checks. OK.
Visible fugitive dust being generated by salt trucks on the property to the south.
2016-04-21 | Performed QC checks beginning at 10:40 CST.
Conducted site checks. OK.
2016-04-25 | Conducted site checks. OK.
2016-04-29 | Conducted site checks. OK.
Cleaned PM10 head.
2016-05-03 | Performed QC checks beginning at 10:10 CST.
Conducted site checks. OK.
Last M = 0.853, ABS = 0.839.
2016-05-05 | Conducted site checks. OK.
2016-05-09 | Conducted site checks. OK.
2016-05-12 | Conducted site checks. OK.
2016-05-16 | Performed QC checks beginning at 10:40 CST.
Conducted site checks. OK.
2016-05-19 | Conducted site checks. OK.
2016-05-24 | Conducted site checks. OK.
2016-05-25 | Elevated concentrations recorded due to high winds. Hour 16:00 experienced high loading.
We believe values reading 986 are true readings representing a concentration of 986 or higher.
2016-05-27 | Performed QC checks beginning at 09:10 CST.
Conducted site checks. OK.
2016-05-31 | Conducted site checks. OK.
2016-06-02 | Conducted site checks. OK.
Cleaned BAM enclosure air conditioner filter.
2016-06-07 | J. Sorkness performed a quarterly audit beginning at 10:40 CST.
Conducted site checks. OK.
2016-06-10 | Conducted site checks. OK.
Changed out BAM filter tape at 07:50 CST,
2016-06-14 | Performed QC checks beginning at 13:40 CST.
Conducted site checks. OK.
2016-06-20 | Cleaned PM10 head.
Conducted site checks. OK.
2016-06-23 | Cleaned enclosure air conditioning filter.
Conducted site checks. OK.
2016-06-27 | Performed QC checks beginning at 09:10 CST.
Conducted site checks. OK.
2016-06-30 | Conducted site checks. OK.
2016-07-06 | Conducted site checks. OK.
2016-07-08 | Cleaned PM10 head.
Conducted site checks. OK.
2016-07-11 | Performed QC checks beginning at 10:40 CST.
Conducted site checks. OK.
2016-07-14 | Conducted site checks. OK.
2016-07-19 | Conducted site checks. OK.
2016-07-22 | Conducted site checks. OK.
2016-07-26 | Performed QC checks beginning at 09:40 CST.
Conducted site checks. OK.
2016-07-28 | Changed out filter tape beginning at 11:50 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
Conducted site checks. OK.
2016-08-02 | Conducted site checks. OK.
2016-08-04 | Power failure during hour 20:00 CST.
2016-08-05 | Conducted site checks. OK.
Found BAM in startup mode, indicating a power failure. Reset to Normal mode.
2016-08-09 | Performed QC checks beginning at 09:10 CST.
Conducted site checks. OK.
2016-08-12 | Conducted site checks. OK.
2016-08-15 | Conducted site checks. OK.
2016-08-18 | Conducted site checks. OK.
2016-08-22 | Performed QC checks beginning at 10:10 CST.
Conducted site checks. OK.
2016-08-24 | Power failure during hour 19:00 CST.
2016-08-26 | Conducted site checks. OK.
Cleaned shelter air condtioner filter.
Found BAM in startup mode, indicating a power failure. Reset to Normal mode.
2016-08-29 | Conducted site checks. OK.
2016-09-01 | J. Sorkness performed a quarterly audit beginning at 09:05 CST.
Conducted site checks. OK.
2016-09-06 | Conducted site checks. OK.
Changed time from 09:53:10 to 09:52 CST.
2016-09-09 | Conducted site checks. OK.
Cleaned PM10 head.
2016-09-13 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
2016-09-16 | Conducted site checks. OK.
2016-09-19 | Conducted site checks. OK.
2016-09-22 | Conducted site checks. OK.
Changed out filter tape starting at 09:40 CST.
2016-10-01 | Conducted site checks. OK.
2016-10-04 | Conducted site checks. OK.
2016-10-07 | Conducted site checks. OK.
2016-10-11 | Performed QC checks beginning at 10:50 CST.
Conducted site checks. OK.
2016-10-14 | Conducted site checks. OK.
Cleaned PM10 sample head.
2016-10-17 | Conducted site checks. OK.
2016-10-20 | Conducted site checks. OK.
2016-10-25 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
Last M = 0.850, ABS = 0.839.
2016-10-28 | Conducted site checks. OK.
2016-11-02 | Conducted site checks. OK.
Salt trucks being loaded to the south generating fugitive dust.
2016-11-04 | Conducted site checks. OK.
Cleaned enclosure AC filter.
2016-11-08 | Conducted site checks. OK.
2016-11-11 | Performed QC checks beginning at 10:35 CST.
Conducted site checks. OK.
2016-11-15 | Conducted site checks. OK.
2016-11-18 | Conducted site checks. OK.
Elevated concentractions due to high southerly winds and salt trucks being loaded on the
property to the south.
2016-11-22 | QC checks were scheduled, but could not be performed because the BAM was
reading a concentration >100 ug/m3.
Conducted site checks. OK.
2016-11-25 | Performed QC checks beginning at 09:20 CST.
Changed out filter tape.
Conducted site checks. OK.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
Conducted site checks. OK.
2016-11-30 | Conducted site checks. OK.
2016-12-02 | J. Sorkness performed a quarterly audit beginning at 09:40 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2016-12-06 | Conducted site checks. OK.
2016-12-09 | Conducted site checks. OK.
2016-12-13 | Performed QC checks beginning at 10:50 CST.
Conducted site checks. OK.
2016-12-16 | Conducted site checks. OK.
Salt trucks being loaded on the property to the south.
2016-12-20 | Conducted site checks. OK.
2016-12-22 | Conducted site checks. OK.
Cleaned PM10 head.
2016-12-28 | Performed QC checks beginning at 13:10 CST.
Conducted site checks. OK.
2016-12-30 | Conducted site checks. OK.
2017-01-03 | Conducted site checks. OK.
2017-01-06 | Conducted site checks. OK.
2017-01-09 | Conducted site checks. OK.
2017-01-13 | Performed QC checks beginning at 10:45 CST.
Conducted site checks. OK.
2017-01-17 | Performed filter tape change out beginning at 11:10 CST.
Conducted site checks. OK.
2017-01-23 | Conducted site checks. OK.
2017-01-27 | Performed QC checks beginning at 10:45 CST.
Conducted site checks. OK.
2017-01-31 | Conducted site checks. OK.
2017-02-06 | Conducted site checks. OK.
Salt barge being unloaded to the south. Fugitive salt observed on top of the enclosure.
2017-02-13 | Performed QC checks beginning at 10:50 CST.
Conducted site checks. OK.
2017-02-17 | Conducted site checks. OK.
2017-02-23 | Conducted site checks. OK.
2017-02-27 | Performed QC checks beginning at 10:50 CST.
Conducted site checks. OK.
2017-03-01 | RAQIS has not been updating from the ST-SW BAM since 19:00 CST on 2/28/17.
A lightning storm had passed through the area that evening. Found the data logger screen
blank. Power cycled the data logger and NL201 without success.
Found water inside the enclosure. Checked out the spare data logger. This data logger
could not get past the initial screen. Sent both data loggers to D. Szatkowski for
for evaluation. The BAM is still collecting valid data.
2017-03-03 | Installed replacement data logger at the ST-SW BAM. Seems to be working properly.
Conducted site checks. OK.
2017-03-08 | Cleaned PM10 heads.
Conducted site checks. OK.
2017-03-16 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
2017-03-20 | Cleaned enclosure AC filters.
Conducted site checks. OK.
Changed time from 09:43:05 to 09:42 CST.
2017-03-24 | J. Sorkness performed a quarterly audit beginning at 09:50 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2017-03-27 | Changed out BAM filter tape beginning at 09:50 CST.
Conducted site checks. OK.
2017-03-31 | Performed QC checks beginning at 08:50 CST.
Conducted site checks. OK.
2017-04-07 | Conducted site checks. OK.
2017-04-10 | Conducted site checks. OK.
2017-04-11 | Power failure from 19:00 to 20:00 CST.
2017-04-13 | Performed QC checks beginning at 09:55 CST.
Conducted site checks. OK.
2017-04-18 | Conducted site checks. OK.
2017-04-21 | Conducted site checks. OK.
2017-04-25 | Performed QC checks beginning at 10:40 CST.
Conducted site checks. OK.
Salt trucks being loaded generating fugitive dust.
2017-04-28 | Conducted site checks. OK.
2017-05-05 | Conducted site checks. OK.
2017-04-25 | Performed QC checks beginning at 10:40 CST.
Conducted site checks. OK.
2017-05-05 | Conducted site checks. OK.
2017-05-12 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
2017-05-16 | Conducted site checks. OK.
2017-05-19 | Conducted site checks. OK.
Cleaned PM10 heads.
2017-05-23 | Conducted site checks. OK.
Changed time from 10:30:10 to 10:29 CST.
2017-05-26 | Performed QC checks beginning at 08:50 CST.
Conducted site checks. OK.
Last M = 0.847, ABS = 0.839.
2017-05-30 | Changed out filter tape beginning at 10:45 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
Leak check as found = 0.3. Leak check as left = 0.5
Conducted site checks. OK.
2017-06-05 | Conducted site checks. OK.
2017-06-08 | Performed QC checks beginning at 09:25 CST.
Conducted site checks. OK.
2017-06-13 | Conducted site checks. OK.
2017-06-20 | Conducted site checks. OK.
2017-06-23 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
Cleaned shelter AC filter.
2017-06-29 | J. Sorkness performed a quarterly audit beginning at 09:05 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2017-07-05 | Conducted site checks. OK.
2017-07-11 | Performed QC checks beginning at 09:45 CST.
Conducted site checks. OK.
2017-07-14 | Conducted site checks. OK.
2017-07-18 | Conducted site checks. OK.
Cleaned PM10 head.
2017-07-21 | Conducted site checks. OK.
2017-07-24 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
2017-07-27 | Conducted site checks. OK.
2017-07-31 | Conducted site checks. OK.
2017-08-04 | Conducted site checks. OK.
Changed out BAM filter tape at 12:50 CST.
2017-08-08 | Performed QC checks beginning at 10:10 CST.
Conducted site checks. OK.
2017-08-09 | Power failure at 12:00 caused by a KCBX employee accidentally turning off a light switch
that provides power to the BAM. BAM resumed normal operation once the switch was
turned back on.
2017-08-11 | Conducted site checks. OK.
2017-08-15 | Conducted site checks. OK.
2017-08-16 | Conducted site checks. OK.
Replaced shelter AC filter.
2017-08-22 | Conducted site checks. OK.
2017-08-25 | Performed QC checks beginning at 09:50 CST.
Changed time from 09:51:05 to 09:50 CST.
Conducted site checks. OK.
2017-08-29 | Conducted site checks. OK.
2017-08-31 | Conducted site checks. OK.
2017-09-05 | Conducted site checks. OK.
Cleaned PM10 head.
2017-09-08 | J. Sorkness performed a quarterly audit beginning at 09:10 CST.
Results indicate all systems to be operating within allowable project limits.
The sampler flow was within audit limits. However, the calibration flow compared to the
design flow was slightly above the allowable limit, so the flow was adjusted.
Target flow As found As left
15.0 15.40 14.52
18.4 19.11 18.13
16.7 17.30 16.33
Conducted site checks. OK.
2017-09-12 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
2017-09-15 | Conducted site checks. OK.
2017-09-19 | Conducted site checks. OK.
2017-09-22 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
2017-09-25 | Investigated the loss of communications. Found the battery backup side of the UPS
no longer holds a charge. Transferred communication equipment power plugs to
other side of the UPS without battery backup.
2017-10-02 | Conducted site checks. OK.
2017-10-05 | Performed QC checks beginning at 10:50 CST.
Conducted site checks. OK.
2017-10-09 | Changed out BAM filter tape at 09:50.
Conducted site checks. OK.
2017-10-12 | Conducted site checks. OK.
2017-10-14 | Hour 2300 was invalidated by the BAM. D. Szatkowski indicated this was caused by
a delta pressure change, possibly by high winds.
2017-10-15 | Hour 0200 was invalidated by the BAM. D. Szatkowski indicated this was caused by
a delta pressure change, possibly by high winds.
2017-10-16 | Conducted site checks. OK.
2017-10-20 | Conducted site checks. OK.
2017-10-24 | Conducted site checks. OK.
2017-10-27 | Performed QC checks beginning at 09:55 CST.
Conducted site checks. OK.
2017-11-06 | Rebuilt the BAM sample pump. Offline at 11:10 CST for the change out.
Conducted site checks. OK.
2017-11-10 | Conducted site checks. OK.
2017-11-14 | Conducted site checks. OK.
2017-11-17 | Performed QC checks beginning at 08:35 CST.
Conducted site checks. OK.
2017-11-24 | Conducted site checks. OK.
2017-11-27 | Performed QC checks beginning at 10:50 CST.
Changed time from 10:49:10 to 10:48 CST.
Conducted site checks. OK.
2017-11-30 | J. Sorkness performed a quarterly audit beginning at 09:35 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2017-12-05 | Conducted site checks. OK.
2017-12-12 | Conducted site checks. OK.
2017-12-15 | Performed QC checks beginning at 10:50 CST.
Conducted site checks. OK.
2017-12-19 | Changed out BAM filter tape beginning at 10:45 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
Leak check as found = 0.4, as left = 0.5.
Conducted site checks. OK.
2017-12-22 | Conducted site checks. OK.
2017-12-26 | Conducted site checks. OK.
2017-12-29 | Conducted site checks. OK.
2018-01-03 | Performed QC checks beginning at 11:50 CST.
Conducted site checks. OK.
2018-01-08 | Power turned off to the site during fuse replacement for the ST-NW BAM.
Hours 09:00 to 11:00 CST lost.
2018-01-09 | Conducted site checks. OK.
2018-01-10 | Power turned off to the site during fuse replacement for the ST-NW BAM.
Hour 13:00 CST lost.
Performed an inspection of the BAM following the power outage.
2018-01-12 | Conducted site checks. OK.
2018-01-17 | Conducted site checks. OK.
2018-01-19 | Performed QC checks beginning at 10:50 CST.
Conducted site checks. OK.
Changed time from 10:48:05 to 10:47 CST.
2018-01-22 | Conducted site checks. OK.
2018-01-25 | Conducted site checks. OK.
2018-01-29 | Performed QC checks beginning at 10:50 CST.
Conducted site checks. OK.
2018-02-02 | Conducted site checks. OK.
2018-02-06 | Conducted site checks. OK.
2018-02-13 | Conducted site checks. OK.
2018-02-16 | Conducted site checks. OK.
2018-02-20 | Performed QC checks beginning at 11:50 CST.
Conducted site checks. OK.
2018-02-23 | Changed out BAM filter tape beginning at 08:45 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
2018-02-26 | Conducted site checks. OK.
2018-03-01 | Conducted site checks. OK.
2018-03-06 | Performed QC checks beginning at 10:50 CST.
Conducted site checks. OK.
2018-03-09 | Conducted site checks. OK.
2018-03-13 | Conducted site checks. OK.
2018-03-16 | Conducted site checks. OK.
2018-03-20 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
2018-03-23 | Conducted site checks. OK.
2018-03-27 | Conducted site checks. OK.
2018-03-29 | J. Sorkness performed a quarterly audit beginning at 09:30 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2018-04-03 | Conducted site checks. OK.
2018-04-06 | Conducted site checks. OK.
2018-04-10 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
2018-04-13 | Changed time from 09:22:10 to 09:21 CST.
Conducted site checks. OK.
2018-04-16 | Conducted site checks. OK.
2018-04-19 | Conducted site checks. OK.
2018-04-23 | Performed QC checks beginning at 09:45 CST.
Conducted site checks. OK.
2018-04-26 | Changed out BAM filter tape beginning at 10:10 CST.
Leak check as found = 0.4, as left = 0.5.
Conducted site checks. OK.
2018-05-01 | Conducted site checks. OK.
2018-05-04 | Conducted site checks. OK.
2018-05-11 | Performed QC checks beginning at 12:50 CST.
Conducted site checks. OK.
2018-05-13 | Hour 07:00 CST invalid due to an instrument malfunction due to a data logger Ethernet
interface failure.
2018-05-18 | Conducted site checks. OK.
2018-05-22 | Conducted site checks. OK.
Cleaned enclosure AC filters.
2018-05-25 | Performed QC checks beginning at 10:10 CST.
Conducted site checks. OK.
D. Szatkowski reported that enclosure temperatures have been at 41 deg C lately
instead of the usual 32 deg C. Cleaned the AC cooling fins with a brush and compressed
air. D. Szatkowski will monitor the situation to see if things improve.
2018-05-29 | Conducted site checks. OK.
2018-06-01 | Conducted site checks. OK.
2018-06-05 | Conducted site checks. OK.
Cleaned enclosure AC filter.
2018-06-08 | Conducted site checks. OK.
2018-06-11 | Performed QC checks beginning at 14:45 CST.
Conducted site checks. OK.
2018-06-14 | The BAM sample pump failed at 14:00 CST on 6/13/2018 and was replaced at 09:00 CST
on 6/14/2018. Data during this period are invalid.
Conducted site checks. OK.
2018-06-18 | Conducted site checks. OK.
2018-06-21 | Conducted site checks. OK.
2018-06-25 | Performed QC checks beginning at 10:10 CST.
Conducted site checks. OK.
2018-06-28 | J. Sorkness performed a quarterly audit beginning at 09:40 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
Cleaned enclosure AC filter.
2018-07-02 | Changed out BAM filter tape beginning at 09:50 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As left leak check = 0.4.
2018-07-05 | Cleaned BAM PM10 head.
Changed time from 10:53:10 to 10:52 CST.
Conducted site checks. OK.
2018-07-10 | Performed QC checks beginning at 08:50 CST.
Conducted site checks. OK.
2018-07-13 | Conducted site checks. OK.
2018-07-16 | Conducted site checks. OK.
2018-07-19 | Conducted site checks. OK.
Moved sample pump away from the data logger to reduce shelter temperature.
2018-07-24 | Conducted site checks. OK.
Checked with D. Szatkowski regarding the recent elevated shelter temperature.
Dave indicated the temperature is now within an allowable range.
2018-07-27 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
2018-07-30 | Conducted site checks. OK.
2018-08-03 | Conducted site checks. OK.
2018-08-07 | Conducted site checks. OK.
2018-08-10 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
2018-08-14 | Conducted site checks. OK.
Cleaned enclosure AC filter.
2018-08-17 | Conducted site checks. OK.
2018-08-23 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2018-08-28 | Conducted site checks. OK.
2018-08-30 | Conducted site checks. OK.
Changed time from 9:50:05 to 9:49 CST.
2018-09-04 | Conducted site checks. OK.
2018-09-07 | Changed out BAM filter tape beginning at 09:45 CST.
Leak check as found = 0.3, as left = 0.5.
Conducted site checks. OK.
2018-09-11 | Conducted site checks. OK.
2018-09-14 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
2018-09-18 | Conducted site checks. OK.
2018-09-21 | Cleaned PM10 sampler head.
Conducted site checks. OK.
2018-09-24 | Performed QC checks beginning at 10:50 CST.
Conducted site checks. OK.
2018-09-27 | J. Sorkness performed a quarterly audit beginning at 10:40 CST, except for PM10 flow.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2018-09-28 | Performed PM10 audit flow checks with DeltaCal s/n 1301 beginning at 09:10 CST.
Results indicate the BAM flows to be within project limits.
2018-10-02 | Conducted site checks. OK.
2018-10-05 | Conducted site checks. OK.
2018-10-12 | Performed QC checks beginning at 08:45 CST.
Conducted site checks. OK.
2018-10-16 | Conducted site checks. OK.
2018-10-19 | Conducted site checks. OK.
2018-10-23 | Conducted site checks. OK.
2018-10-26 | Performed QC checks beginning at 09:10 CST.
Changed time from 09:08:05 to 09:07 CST.
Conducted site checks. OK.
2018-10-30 | Conducted site checks. OK.
Cleaned enclosure AC filter.
2018-11-02 | Conducted site checks. OK.
2018-11-06 | Conducted site checks. OK.
2018-11-12 | Changed out BAM filter tape beginning at 10:50 CST.
Leak check as found = 0.6, as left = 0.5.
Conducted site checks. OK.
2018-11-16 | Performed QC checks beginning at 10:50 CST.
Conducted site checks. OK.
2018-11-27 | Conducted site checks. OK.
2018-11-30 | Performed QC checks beginning at 10:50 CST.
Conducted site checks. OK.
2018-12-03 | Conducted site checks. OK.
2018-12-06 | J. Sorkness performed a quarterly audit beginning at 09:45 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2018-12-11 | Conducted site checks. OK.
2018-12-14 | Performed QC checks beginning at 10:50 CST.
Conducted site checks. OK.
2018-12-18 | Changed time from 11:09:05 to 11:08 CST.
Conducted site checks. OK.
2018-12-21 | Conducted site checks. OK.
Cleaned PM10 head.
2018-12-28 | Performed QC checks beginning at 10:10 CST.
Conducted site checks. OK.
2018-12-31 | Conducted site checks. OK.
2019-01-08 | Conducted site checks. OK.
2019-01-10 | Power failure during hour 11:00 CST.
2019-01-11 | Performed QC checks beginning at 10:50 CST.
Conducted site checks. OK.
2019-01-15 | Conducted site checks. OK.
2019-01-18 | Changed out BAM filter tape beginning at 10:50 CST.
Leak check as found = 0.6, as left = 0.5.
Conducted site checks. OK.
2019-01-24 | Conducted site checks. OK.
2019-02-01 | Conducted site checks. OK.
2019-02-07 | Performed QC checks beginning at 13:50 CST.
Conducted site checks. OK.
2019-02-11 | Conducted site checks. OK.
Changed time from 10:32:05 to 10:31 CST.
2019-02-15 | Conducted site checks. OK.
2019-02-21 | Performed QC checks beginning at 14:35 CST.
Conducted site checks. OK.
2019-02-25 | Conducted site checks. OK.
2019-03-01 | Conducted site checks. OK.
2019-03-06 | Performed QC checks beginning at 12:50 CST.
Conducted site checks. OK.
2019-03-12 | Conducted site checks. OK.
2019-03-15 | Trimmed the end of the sample pump line at 09:45 CST.
Conducted site checks. OK.
2019-03-18 | Performed a leak check at 10:10 CST.
As found leak check = 0.5.
Conducted site checks. OK.
2019-03-20 | M. Wolfskill performed a quarterly audit beginning at 09:20 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2019-03-25 | Performed QC checks beginning at 10:30 CST.
Changed out BAM filter tape beginning at 09:10 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As found leak check = 0.5, As left leak check = 0.5.
2019-03-29 | Conducted site checks. OK.
2019-04-02 | Conducted site checks. OK.
Cleaned PM10 head.
Changed time from 10:17:05 to 10:16 CST.
2019-04-05 | Conducted site checks. OK.
2019-04-09 | Conducted site checks. OK.
2019-04-12 | Performed QC checks beginning at 09:45 CST.
Conducted site checks. OK.
2019-04-16 | Conducted site checks. OK.
Inspected enclosure HVAC filter. Found it clean.
2019-04-18 | Conducted site checks. OK.
2019-04-23 | Conducted site checks. OK.
2019-04-26 | Conducted site checks. OK.
2019-04-30 | Performed QC checks beginning at 09:45 CST.
Conducted site checks. OK.
2019-05-03 | Conducted site checks. OK.
2019-05-06 | Conducted site checks. OK.
2019-05-10 | Conducted site checks. OK.
2019-05-16 | Performed QC checks beginning at 10:55 CST.
Conducted site checks. OK.
2019-05-22 | Conducted site checks. OK.
2019-05-24 | Conducted site checks. OK.
2019-05-28 | Performed QC checks beginning at 11:45 CST.
Changed time from 12:06:05 to 12:05 CST.
Changed out BAM filter tape.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As found leak check = 0.4, As left leak check = 0.5.
Conducted site checks. OK.
2019-06-04 | Cleaned PM10 sampler head.
Conducted site checks. OK.
2019-06-11 | Performed QC checks beginning at 09:45 CST.
Conducted site checks. OK.
2019-06-14 | Conducted site checks. OK.
2019-06-18 | Cleaned shelter HVAC filter.
Conducted site checks. OK.
2019-06-21 | Conducted site checks. OK.
2019-06-26 | M. Wolfskill performed a quarterly audit beginning at 10:05 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2019-06-28 | Performed QC checks beginning at 08:50 CST.
Conducted site checks. OK.
2019-07-02 | Conducted site checks. OK.
2019-07-05 | Cleaned shelter HVAC filter.
Conducted site checks. OK.
2019-07-09 | Conducted site checks. OK.
2019-07-12 | Performed QC checks beginning at 08:50 CST.
Conducted site checks. OK.
2019-07-16 | Power failure during hour 12:00 CST.
Conducted site checks. OK.
2019-07-19 | Conducted site checks. OK.
2019-07-23 | Cleaned PM10 sampler head.
Power failure during hours 09:00 to 11:00 CST. KCBX electricians were doing maintenance.
Conducted site checks. OK.
2019-07-26 | Performed QC checks beginning at 09:50 CST.
Changed time from 09:48:05 to 09:47 CST.
Conducted site checks. OK.
2019-07-30 | Changed out BAM filter tape at 09:45 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As found leak check = 0.4, As left leak check = 0.5.
Conducted site checks. OK.
2019-08-02 | Conducted site checks. OK.
2019-08-06 | Conducted site checks. OK.
2019-08-09 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
2019-08-13 | Conducted site checks. OK.
2019-08-16 | Conducted site checks. OK.
2019-08-20 | Cleaned shelter HVAC filter.
Conducted site checks. OK.
2019-08-23 | Performed QC checks beginning at 09:40 CST.
Conducted site checks. OK.
2019-08-27 | Conducted site checks. OK.
2019-09-03 | Conducted site checks. OK.
2019-09-05 | Conducted site checks. OK.
2019-09-10 | Performed QC checks beginning at 09:45 CST.
Conducted site checks. OK.
2019-09-11 | Hours 03:00 and 04:00 CST are invalid due to a power failure.
2019-09-13 | Conducted site checks. OK.
2019-09-19 | M. Wolfskill performed a quarterly audit beginning at 08:10 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2019-09-24 | Conducted site checks. OK.
2019-09-27 | Performed QC checks beginning at 09:40 CST.
Changed time from 09:40:05 to 09:39 CST.
Conducted site checks. OK.
2019-09-30 | Conducted site checks. OK.
2019-10-04 | Changed out BAM filter tape at 09:30 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As found leak check = 0.5, As left leak check = 0.5.
Conducted site checks. OK.
2019-10-07 | Conducted site checks. OK.
2019-10-09 | Conducted site checks. OK.
2019-10-15 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
2019-10-21 | Conducted site checks. OK.
2019-10-24 | Conducted site checks. OK.
2019-10-28 | Performed QC checks beginning at 11:50 CST.
Conducted site checks. OK.

## ST-SE
2014-10-07 | Performed BAM QC checks starting at 13:00 CST.
Cleaned PM10 sample.
Leak check as found was 0.3, so I did not clean the nozzle.
Conducted site checks. OK.
2014-10-10 | Conducted site checks. OK.
2014-10-13 | Conducted site checks. OK.
Found BAM not in "Normal Mode". There was a lost hour of data overnight. The BAM
seemed to be operating correctly. Placed the BAM in "Normal Mode".
2014-10-16 | Conducted site checks. OK.
2014-10-21 | Performed BAM QC checks starting at 10:40 CST.
Conducted site checks. OK.
2014-10-24 | Conducted site checks. OK.
2014-10-28 | Conducted site checks. OK.
2014-10-30 | Conducted site checks. OK.
2014-11-03 | Performed BAM QC checks starting at 13:30 CST.
Changed time from 13:30:05 to 13:32 CST.
Conducted site checks. OK.
2014-11-05 | Conducted site checks. OK.
2014-11-10 | Conducted site checks. OK.
2014-11-14 | Conducted site checks. OK.
2014-11-17 | Conducted site checks. OK.
2014-11-21 | Performed BAM QC checks starting at 11:40 CST.
Replaced filter tape.
Conducted site checks. OK.
2014-11-26 | Conducted site checks. OK.
2014-12-01 | Conducted site checks. OK.
2014-12-05 | Performed BAM QC checks starting at 12:10 CST.
Conducted site checks. OK.
Cleaned PM10 sample head.
2014-12-08 | J. Sorkness conducted audits of the BAM.
Conducted site checks. OK.
2014-12-12 | Conducted site checks. OK.
2014-12-16 | Conducted site checks. OK.
2014-12-18 | Performed BAM QC checks starting at 11:00 CST.
Conducted site checks. OK.
2014-12-23 | Conducted site checks. OK.
2014-12-30 | Conducted site checks. OK.
Changed time from 10:02:05 to 10:01 CST.
2015-01-02 | Conducted site checks. OK.
2015-01-05 | Conducted site checks. OK.
2015-01-08 | Conducted site checks. OK.
2015-01-12 | Perform QC checks beginning at 12:00 CST.
Conducted site checks. OK.
2015-01-13 | Begin annual zeroing at 10:15 CST.
Conducted site checks. OK.
Changed units' background concentration to 0.0000.
2015-01-16 | Conducted site checks. OK.
Concluded annual zeroing at 11:40 CST.
Changed units' background concentration to -0.0050.
2015-01-20 | Input new zero background concentration of -0.0066 into BAM at 10:57 CST.
Conducted site checks. OK.
2015-01-23 | Conducted site checks. OK.
2015-01-26 | Perform QC checks beginning at 11:20 CST.
Conducted site checks. OK.
2015-01-29 | Conducted site checks. OK.
2015-02-04 | Conducted site checks. OK.
2015-02-06 | Conducted site checks. OK.
2015-02-10 | Perform QC checks beginning at 12:50 CST.
Conducted site checks. OK.
2015-02-13 | Conducted site checks. OK.
2015-02-16 | Conducted site checks. OK.
Changed time from 09:32:15 to 09:31 CST.
2015-02-19 | Conducted site checks. OK.
2015-02-24 | Conducted site checks. OK.
2015-02-27 | Perform QC checks beginning at 12:50 CST.
Conducted site checks. OK.
2015-03-02 | Conducted site checks. OK.
2015-03-06 | Changed filter tape beginning at 13:50 CST.
Performed leak check = 0.5.
Cleaned PM10 head.
Conducted site checks. OK.
2015-03-09 | Conducted site checks. OK.
2015-03-11 | Conducted site checks. OK.
2015-03-17 | Conducted site checks. OK.
2015-03-19 | J. Sorkness performed audits beginning at 12:30 CST.
2015-03-20 | Conducted site checks. OK.
2015-03-24 | Perform QC checks beginning at 11:10 CST.
Conducted site checks. OK.
2015-03-27 | Conducted site checks. OK.
2015-03-31 | Perform leak checks beginning at 11:45 CST.
Leak as found = 0.5, as left = 0.5.
Conducted site checks. OK.
2015-04-01 | Conducted site checks. OK.
2015-04-08 | Perform QC checks beginning at 12:00 CST.
Conducted site checks. OK.
2015-04-10 | Conducted site checks. OK.
2015-04-15 | Conducted site checks. OK.
Changed time from 12:32:05 to 12:31 CST.
2015-04-17 | Conducted site checks. OK.
2015-04-24 | Perform QC checks beginning at 11:00 CST.
Conducted site checks. OK.
2015-04-27 | Conducted site checks. OK.
Installed new alarm wiring beginning at 12:20 CST.
2015-04-30 | Conducted site checks. OK.
Changed filter tape beginning at 12:15 CST.
2015-05-04 | Perform QC checks beginning at 10:50 CST.
Conducted site checks. OK.
2015-05-07 | Conducted site checks. OK.
2015-05-11 | Conducted site checks. OK.
2015-05-14 | Conducted site checks. OK.
2015-05-18 | Conducted site checks. OK.
2015-05-21 | Perform QC checks beginning at 11:40 CST.
Conducted site checks. OK.
2015-05-27 | Conducted site checks. OK.
2015-05-29 | Conducted site checks. OK.
2015-06-01 | Conducted site checks. OK.
2015-06-04 | Perform QC checks beginning at 12:40 CST.
Changed time from 12:40:00 to 12:39 CST.
Cleaned nozzle and screen.
Cleaned PM10 head.
Conducted site checks. OK.
2015-06-09 | Conducted site checks. OK.
2015-06-12 | Conducted site checks. OK.
2015-06-15 | Performed leak check beginning at 12:35 CST.
Leak check = 0.3.
Conducted site checks. OK.
2015-06-17 | J. Sorkness performed audit checks beginning at 07:50 CST.
2015-06-18 | Conducted site checks. OK.
2015-06-22 | Conducted site checks. OK.
2015-06-24 | Conducted site checks. OK.
Changed filter tape at 08:10 CST.
2015-06-30 | Conducted site checks. OK.
2015-07-02 | Performed QC checks beginning at 08:55 CST.
Conducted site checks. OK.
2015-07-07 | Conducted site checks. OK.
2015-07-10 | Conducted site checks. OK.
2015-07-14 | Performed QC checks beginning at 09:40 CST.
Conducted site checks. OK.
2015-07-16 | Conducted site checks. OK.
2015-07-22 | Conducted site checks. OK.
2015-07-24 | Conducted site checks. OK.
2015-07-28 | Performed QC checks beginning at 11:05 CST.
Changed time from 11:09:05 to 11:08 CST.
Conducted site checks. OK.
2015-07-30 | Conducted site checks. OK.
2015-08-05 | Conducted site checks. OK.
Weed control contractor spraying near and on shelter.
2015-08-07 | Conducted site checks. OK.
2015-08-11 | Performed a BAM leak check beginning at 11:30 CST.
Conducted site checks. OK.
2015-08-12 | J. Sorkness performed a quarterly audit beginning at 08:50 CST.
Conducted site checks. OK.
2015-08-14 | Conducted site checks. OK.
2015-08-18 | Power failure caused by a lightning strike at 19:00 CST.
2015-08-19 | Power restored. Valid data collection resumed at 17:00 CST.
2015-08-21 | Performed QC checks beginning at 11:15 CST.
Changed BAM filter tape.
Conducted site checks. OK.
2015-08-28 | Conducted site checks. OK.
Took BAM offline to record Last M and ABS readings at 10:30 CST.
Last M = 0.788, ABS = 0.778
2015-09-01 | Conducted site checks. OK.
2015-09-04 | Performed QC checks beginning at 09:50 CST.
Adjusted temperature from 25.8 to 27.6 degrees C.
Conducted site checks. OK.
2015-09-09 | Conducted site checks. OK.
2015-09-11 | Conducted site checks. OK.
Cleaned PM10 head.
2015-09-15 | Conducted site checks. OK.
2015-09-16 | Replaced the sample pump with a rebuilt one. Hours 10:00 and 11:00 CST are invalid due
to the change out.
Conducted site checks. OK.
2015-09-22 | Performed QC checks beginning at 10:20 CST.
Began 72-hour zero calibration following QC checks.
Changed zero offset to 0.0000.
Conducted site checks. OK.
2015-09-25 | Concluded zero calibration.
Adjusted zero offset to former number of -0.0066.
New zero offset numbers will be input next week.
Valid data collection resumed at 12:00 CST.
Conducted site checks. OK.
2015-09-28 | Placed new zero offset into monitor at 11:15 CST.
The new zero offset is -0.0028.
Conducted site checks. OK.
2015-09-30 | Conducted site checks. OK.
2015-10-02 | Conducted site checks. OK.
2015-10-05 | Performed QC checks beginning at 12:35 CST.
Conducted site checks. OK.
2015-10-06 | KCBX power failure from 02:00 to 03:00 CST.
2015-10-08 | Conducted site checks. OK.
2015-10-13 | Conducted site checks. OK.
2015-10-15 | Conducted site checks. OK.
2015-10-19 | Conducted site checks. OK.
2015-10-22 | Performed QC checks beginning at 13:35 CST.
Changed filter tape.
Conducted site checks. OK.
2015-10-27 | Conducted site checks. OK.
2015-10-29 | Conducted site checks. OK.
2015-11-02 | Conducted site checks. OK.
2015-11-04 | Performed QC checks beginning at 15:20 CST.
Changed time from 15:21:00 to 15:20 ST.
Conducted site checks. OK.
2015-11-10 | Conducted site checks. OK.
Changed time from 11:37:00 to 11:36 CST.
2015-11-12 | Conducted site checks. OK.
2015-11-16 | Performed a leak check beginning at 13:10 CST.
Leak = 0.5.
Conducted site checks. OK.
2015-11-18 | J. Sorkness performed a quarterly audit beginning at 09:35 CST.
Conducted site checks. OK.
2015-11-23 | Conducted site checks. OK.
2015-11-25 | Conducted site checks. OK.
2015-11-30 | Performed QC checks beginning at 13:20 CST.
Conducted site checks. OK.
2015-12-03 | Conducted site checks. OK.
2015-12-08 | Conducted site checks. OK.
2015-12-10 | Conducted site checks. OK.
Cleaned PM10 head.
2015-12-14 | Conducted site checks. OK.
2015-12-16 | Performed QC checks beginning at 12:20 CST.
Conducted site checks. OK.
2015-12-22 | Changed filter tape beginning at 11:55 CST.
Conducted site checks. OK.
2015-12-29 | Performed QC checks beginning at 10:50 CST.
Conducted site checks. OK.
2015-12-31 | Conducted site checks. OK.
2016-01-04 | Conducted site checks. OK.
2016-01-06 | Conducted site checks. OK.
2016-01-11 | Conducted site checks. OK.
Changed time from 11:11:10 to 11:10 CST.
2016-01-14 | Conducted site checks. OK.
2016-01-19 | Performed QC checks beginning at 11:35 CST.
Conducted site checks. OK.
2016-01-21 | Conducted site checks. OK.
2016-01-25 | Conducted site checks. OK.
2016-01-27 | Conducted site checks. OK.
2016-02-01 | Conducted site checks. OK.
2016-02-04 | Conducted site checks. OK.
2016-02-08 | Performed QC checks beginning at 10:40 CST.
Changed filter tape.
Conducted site checks. OK.
2016-02-11 | Conducted site checks. OK.
2016-02-16 | Conducted site checks. OK.
2016-02-19 | Conducted site checks. OK.
2016-02-23 | Performed QC checks beginning at 10:45 CST.
Changed time from 10:45:10 to 10:44 CST.
Conducted site checks. OK.
2016-02-26 | Conducted site checks. OK.
2016-02-29 | Conducted site checks. OK.
2016-03-04 | Conducted site checks. OK.
2016-03-08 | Performed QC checks beginning at 10:35 CST.
Conducted site checks. OK.
2016-03-11 | Conducted site checks. OK.
Cleaned PM10 head.
2016-03-14 | Conducted site checks. OK.
2016-03-17 | Conducted site checks. OK.
Cleaned enclosure air conditioner and changed filter.
2016-03-23 | J. Sorkness performed a quarterly audit beginning at 11:10 CST.
Conducted site checks. OK.
2016-03-25 | Conducted site checks. OK.
2016-04-01 | Conducted site checks. OK.
2016-04-03 | Power failure at 18:00 CST.
2016-04-04 | Performed QC checks beginning at 10:40 CST.
Conducted site checks. OK.
2016-04-07 | Conducted site checks. OK.
Changed time from 10:48:05 to 10:47 CST.
2016-04-11 | Performed preventative maintenace beginning at 11:25 CST.
Cleaned nozzle, screen, and tape rollers.
Changed filter tape.
Conducted site checks. OK.
2016-04-14 | Conducted site checks. OK.
2016-04-18 | Conducted site checks. OK.
2016-04-21 | Performed QC checks beginning at 10:10 CST.
Conducted site checks. OK.
2016-04-25 | Conducted site checks. OK.
2016-04-29 | Conducted site checks. OK.
Cleaned PM10 head.
2016-05-03 | Performed QC checks beginning at 10:40 CST.
Conducted site checks. OK.
Last M = 0.794 ABS = 0.778.
2016-05-05 | Conducted site checks. OK.
2016-05-09 | Conducted site checks. OK.
2016-05-12 | Conducted site checks. OK.
2016-05-16 | Performed QC checks beginning at 11:20 CST.
Conducted site checks. OK.
2016-05-19 | Conducted site checks. OK.
2016-05-24 | Conducted site checks. OK.
2016-05-25 | Elevated concentrations recorded due to high winds. Hour 16:00 experienced high loading.
We believe values reading 986 are true readings representing a concentration of 986 or higher.
2016-05-27 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2016-05-31 | Conducted site checks. OK.
2016-06-02 | Conducted site checks. OK.
Cleaned BAM enclosure air conditioner filter.
Changed time from 10:46:05 to 10:45 CST.
2016-06-07 | J. Sorkness performed a quarterly audit beginning at 10:50 CST.
Conducted site checks. OK.
2016-06-10 | Conducted site checks. OK.
Changed out BAM filter tape at 07:30 CST,
2016-06-14 | Performed QC checks beginning at 14:25 CST.
Conducted site checks. OK.
2016-06-20 | Cleaned PM10 head.
Conducted site checks. OK.
2016-06-23 | Cleaned enclosure air conditioning filter.
Conducted site checks. OK.
2016-06-27 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2016-06-30 | Conducted site checks. OK.
2016-07-06 | Conducted site checks. OK.
2016-07-08 | Cleaned PM10 head.
Conducted site checks. OK.
2016-07-11 | Performed QC checks beginning at 11:10 CST.
Conducted site checks. OK.
2016-07-14 | Conducted site checks. OK.
2016-07-19 | Conducted site checks. OK.
2016-07-22 | Conducted site checks. OK.
2016-07-26 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
2016-07-28 | Changed out filter tape beginning at 12:10 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
Conducted site checks. OK.
2016-08-02 | Conducted site checks. OK.
Changed time from 10:10:05 to 10:09 CST.
2016-08-05 | Conducted site checks. OK.
2016-08-09 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2016-08-12 | Conducted site checks. OK.
2016-08-15 | Conducted site checks. OK.
2016-08-18 | Conducted site checks. OK.
2016-08-22 | Performed QC checks beginning at 10:30 CST.
Conducted site checks. OK.
2016-08-26 | Conducted site checks. OK.
Cleaned shelter air condtioner filter.
2016-08-29 | Conducted site checks. OK.
2016-09-01 | J. Sorkness performed a quarterly audit beginning at 09:20 CST.
Conducted site checks. OK.
2016-09-06 | Conducted site checks. OK.
2016-09-09 | Conducted site checks. OK.
Cleaned PM10 head.
2016-09-13 | Performed QC checks beginning at 10:25 CST.
Conducted site checks. OK.
2016-09-16 | Conducted site checks. OK.
2016-09-19 | Conducted site checks. OK.
2016-09-22 | Conducted site checks. OK.
Changed out filter tape starting at 09:50 CST.
2016-10-01 | Conducted site checks. OK.
Changed time from 10:16:10 to 10:15 CST.
2016-10-04 | Conducted site checks. OK.
2016-10-07 | Conducted site checks. OK.
2016-10-11 | Performed QC checks beginning at 11:15 CST.
Conducted site checks. OK.
2016-10-14 | Conducted site checks. OK.
Cleaned PM10 sample head.
2016-10-17 | Conducted site checks. OK.
2016-10-20 | Conducted site checks. OK.
2016-10-25 | Performed QC checks beginning at 10:20 CST.
Conducted site checks. OK.
Last M = 0.799, ABS = 0.778.
2016-10-28 | Conducted site checks. OK.
2016-11-02 | Conducted site checks. OK.
2016-11-04 | Conducted site checks. OK.
Cleaned enclosure AC filter.
2016-11-08 | Conducted site checks. OK.
2016-11-11 | Performed QC checks beginning at 10:20 CST.
Conducted site checks. OK.
2016-11-15 | Conducted site checks. OK.
Changed time from 09:41:05 to 09:40 CST.
2016-11-18 | Conducted site checks. OK.
Elevated concentractions due to high southerly winds and salt trucks being loaded on the
property to the south.
2016-11-22 | Performed QC checks beginning at 11:20 CST.
Changed out filter tape.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
Conducted site checks. OK.
2016-11-25 | Conducted site checks. OK.
2016-11-30 | Conducted site checks. OK.
2016-12-02 | J. Sorkness performed a quarterly audit beginning at 08:40 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2016-12-06 | Conducted site checks. OK.
2016-12-09 | Conducted site checks. OK.
2016-12-13 | Performed QC checks beginning at 11:10 CST.
Conducted site checks. OK.
2016-12-16 | Conducted site checks. OK.
2016-12-20 | Conducted site checks. OK.
2016-12-22 | Conducted site checks. OK.
Cleaned PM10 head.
2016-12-28 | Performed QC checks beginning at 13:35 CST.
Conducted site checks. OK.
2016-12-30 | Conducted site checks. OK.
2017-01-03 | Conducted site checks. OK.
2017-01-06 | Conducted site checks. OK.
Changed time from 14:34:05 to 14:33.
2017-01-09 | Conducted site checks. OK.
2017-01-13 | Performed QC checks beginning at 11:10 CST.
Conducted site checks. OK.
2017-01-17 | Performed filter tape change out beginning at 11:25 CST.
Conducted site checks. OK.
2017-01-23 | Conducted site checks. OK.
2017-01-27 | Performed QC checks beginning at 11:15 CST.
Conducted site checks. OK.
2017-01-31 | Conducted site checks. OK.
2017-02-06 | Conducted site checks. OK.
2017-02-13 | Performed QC checks beginning at 11:15 CST.
Conducted site checks. OK.
2017-02-17 | Conducted site checks. OK.
2017-02-23 | Conducted site checks. OK.
2017-02-27 | Performed QC checks beginning at 10:35 CST.
Changed time from 10:33:05 to 10:32 CST.
Conducted site checks. OK.
2017-03-03 | Conducted site checks. OK.
2017-03-08 | Cleaned PM10 heads.
Conducted site checks. OK.
2017-03-16 | Performed QC checks beginning at 09:40 CST.
Conducted site checks. OK.
2017-03-20 | Cleaned enclosure AC filters.
Conducted site checks. OK.
2017-03-24 | J. Sorkness performed a quarterly audit beginning at 09:50 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2017-03-27 | Changed out BAM filter tape beginning at 09:40 CST.
Conducted site checks. OK.
2017-03-31 | Performed QC checks beginning at 08:35 CST.
Conducted site checks. OK.
2017-04-07 | Conducted site checks. OK.
2017-04-10 | Conducted site checks. OK.
2017-04-13 | Performed QC checks beginning at 09:45 CST.
Conducted site checks. OK.
2017-04-18 | Power failure from 07:00 to 08:00 CST.
Conducted site checks. OK.
2017-04-21 | Conducted site checks. OK.
Changed time from 12:22:10 to 12:21 CST.
2017-04-25 | Performed QC checks beginning at 09:25 CST.
Conducted site checks. OK.
2017-04-28 | Conducted site checks. OK.
2017-05-05 | Conducted site checks. OK.
2017-05-12 | Performed QC checks beginning at 09:35 CST.
Conducted site checks. OK.
2017-05-16 | Conducted site checks. OK.
2017-05-19 | Conducted site checks. OK.
Cleaned PM10 heads.
2017-05-23 | Conducted site checks. OK.
2017-05-26 | Performed QC checks beginning at 09:15 CST.
Conducted site checks. OK.
Last M = 0.800, ABS = 0.778.
2017-05-30 | Changed out filter tape beginning at 10:25 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
Leak check as found = 0.5. Leak check as left = 0.5
Conducted site checks. OK.
2017-06-05 | Conducted site checks. OK.
2017-06-08 | Performed QC checks beginning at 09:40 CST.
Conducted site checks. OK.
2017-06-13 | Conducted site checks. OK.
Changed time from 09:48:05 to 09:47 CST.
2017-06-14 | A storm in the area blew out the fuse to the ST-SE BAM at 12:00 CST.
2017-06-15 | The fuse was replaced at 11:20 CST by the local power company. Valid data collection
resumed.
2017-06-20 | Conducted site checks. OK.
2017-06-23 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
Cleaned shelter AC filter.
2017-06-29 | J. Sorkness performed a quarterly audit beginning at 09:30 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2017-07-05 | Conducted site checks. OK.
2017-07-11 | Performed QC checks beginning at 10:15 CST.
Conducted site checks. OK.
2017-07-14 | Conducted site checks. OK.
2017-07-18 | Conducted site checks. OK.
Cleaned PM10 head.
2017-07-21 | Conducted site checks. OK.
2017-07-24 | Performed QC checks beginning at 09:40 CST.
Conducted site checks. OK.
2017-07-27 | Conducted site checks. OK.
2017-07-31 | Conducted site checks. OK.
Changed time from 09:29:05 to 09:28 CST.
2017-08-04 | Conducted site checks. OK.
Changed out BAM filter tape at 12:30 CST.
2017-08-08 | Performed QC checks beginning at 09:45 CST.
Conducted site checks. OK.
2017-08-11 | Conducted site checks. OK.
2017-08-15 | Conducted site checks. OK.
2017-08-16 | Conducted site checks. OK.
Replaced shelter AC filter.
2017-08-22 | Conducted site checks. OK.
2017-08-25 | Performed QC checks beginning at 09:35 CST.
Conducted site checks. OK.
2017-08-29 | Conducted site checks. OK.
2017-08-31 | Conducted site checks. OK.
2017-09-05 | Conducted site checks. OK.
Cleaned PM10 head.
2017-09-08 | J. Sorkness performed a quarterly audit beginning at 09:20 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2017-09-12 | Performed QC checks beginning at 10:15 CST.
Conducted site checks. OK.
2017-09-15 | Conducted site checks. OK.
2017-09-19 | Conducted site checks. OK.
2017-09-22 | Performed QC checks beginning at 09:40 CST.
Conducted site checks. OK.
2017-10-02 | Conducted site checks. OK.
2017-10-05 | Performed QC checks beginning at 10:35 CST.
Conducted site checks. OK.
2017-10-09 | Changed out BAM filter tape at 10:10.
Conducted site checks. OK.
2017-10-12 | Conducted site checks. OK.
2017-10-16 | Conducted site checks. OK.
2017-10-20 | Conducted site checks. OK.
2017-10-24 | Conducted site checks. OK.
2017-10-27 | Performed QC checks beginning at 09:40 CST.
Conducted site checks. OK.
2017-11-06 | Rebuilt the BAM sample pump. Offline at 12:10 CST for the change out.
Conducted site checks. OK.
2017-11-10 | Conducted site checks. OK.
2017-11-14 | Conducted site checks. OK.
Changed time from 10:47:10 to 10:46 CST.
2017-11-17 | Performed QC checks beginning at 08:45 CST.
Conducted site checks. OK.
2017-11-24 | Conducted site checks. OK.
2017-11-27 | Performed QC checks beginning at 11:10 CST.
Conducted site checks. OK.
2017-11-30 | J. Sorkness performed a quarterly audit beginning at 09:45 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2017-12-05 | Conducted site checks. OK.
2017-12-12 | Conducted site checks. OK.
2017-12-15 | Performed QC checks beginning at 11:15 CST.
Conducted site checks. OK.
2017-12-19 | Changed out BAM filter tape beginning at 11:15 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
Leak check as found = 0.6, as left = 0.5.
Conducted site checks. OK.
2017-12-22 | Conducted site checks. OK.
2017-12-26 | Conducted site checks. OK.
2017-12-29 | Conducted site checks. OK.
Changed time from 09:34:05 to 09:33 CST.
2018-01-03 | Performed QC checks beginning at 11:40 CST.
Conducted site checks. OK.
2018-01-09 | Conducted site checks. OK.
2018-01-12 | Conducted site checks. OK.
Changed time from 10:40:05 to 10:39 CST.
2018-01-17 | Conducted site checks. OK.
2018-01-19 | Performed QC checks beginning at 10:30 CST.
Conducted site checks. OK.
2018-01-22 | Conducted site checks. OK.
2018-01-25 | Conducted site checks. OK.
2018-01-29 | Performed QC checks beginning at 11:15 CST.
Conducted site checks. OK.
2018-02-02 | Conducted site checks. OK.
Seems like regular heater is not working properly. Will add an auxiliary heater on next
visit.
2018-02-06 | Conducted site checks. OK.
Installed auxiliary heater.
2018-02-07 | Auxiliary heater tripped the circuit breaker. Took out of service.
Hours 2/6/18 16:00 to 2/7/18 09:00 CST lost.
2018-02-13 | Conducted site checks. OK.
2018-02-16 | Conducted site checks. OK.
Found 1/8" rip on each side of the tape past the rollers. Cleaned rollers and capstan and
advanced the tape. BAM taken offline at 10:00 CST.
2018-02-20 | Performed QC checks beginning at 11:40 CST.
Conducted site checks. OK.
Found the same tape rips as on 2/16. Advanced the tape.
2018-02-23 | Found filter tape had broken. Changed out BAM filter tape beginning at 08:45 CST.
Hours 2/22/18 20:00 to 2/23/18 08:00 CST lost due to the tape break.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
2018-02-26 | Conducted site checks. OK.
The new tape seems to be working fine. With the help of KCBX staff, opened the HVAC
door. Moved the max/min adjustment on the thermostat back and forth several times
in case there were contact issues. D. Szatkowski advised that the temperature inside
the enclosure is OK.
2018-03-01 | Conducted site checks. OK.
2018-03-06 | Performed QC checks beginning at 11:15 CST.
Conducted site checks. OK.
2018-03-09 | Conducted site checks. OK.
2018-03-13 | Conducted site checks. OK.
2018-03-16 | Conducted site checks. OK.
2018-03-20 | Performed QC checks beginning at 11:10 CST.
Conducted site checks. OK.
2018-03-23 | Conducted site checks. OK.
2018-03-27 | Conducted site checks. OK.
2018-03-29 | J. Sorkness performed a quarterly audit beginning at 09:50 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2018-04-03 | Conducted site checks. OK.
2018-04-06 | Conducted site checks. OK.
2018-04-10 | Performed QC checks beginning at 10:15 CST.
Conducted site checks. OK.
2018-04-13 | Conducted site checks. OK.
2018-04-16 | Conducted site checks. OK.
2018-04-19 | Conducted site checks. OK.
2018-04-23 | Unable to perform QC checks due to high recorded concentration (218 ug/m3).
Observed dump trucks next to the site generating fugitive dust.
Conducted site checks. OK.
2018-04-26 | Performed QC checks beginning at 10:25 CST.
Changed out BAM filter tape beginning at 10:40 CST.
Leak check as found = 0.6, as left = 0.5.
Conducted site checks. OK.
2018-05-01 | Conducted site checks. OK.
2018-05-04 | Conducted site checks. OK.
2018-05-11 | Performed QC checks beginning at 13:10 CST.
Conducted site checks. OK.
2018-05-18 | Conducted site checks. OK.
2018-05-22 | Conducted site checks. OK.
Cleaned enclosure AC filters.
2018-05-25 | Performed QC checks beginning at 09:45 CST.
Conducted site checks. OK.
Comms to this site stopped at 22:00 CST on 5/24/18. Power cycled the comm equipment.
The comm system came back online.
Hour 06:00 is invalid due to KCBX personnel power cycling everything, including the BAM.
D. Szatkowski reported that enclosure temperatures have been at 41 deg C lately
instead of the usual 32 deg C. Cleaned the AC cooling fins with a brush and compressed
air. D. Szatkowski will monitor the situation to see if things improve.
2018-05-29 | Conducted site checks. OK.
2018-06-01 | Conducted site checks. OK.
2018-06-05 | Conducted site checks. OK.
Cleaned enclosure AC filter.
2018-06-08 | Conducted site checks. OK.
2018-06-11 | Performed QC checks beginning at 13:50 CST.
Conducted site checks. OK.
2018-06-14 | Conducted site checks. OK.
2018-06-18 | Conducted site checks. OK.
2018-06-21 | Conducted site checks. OK.
2018-06-25 | Performed QC checks beginning at 10:20 CST.
Conducted site checks. OK.
2018-06-28 | J. Sorkness performed a quarterly audit beginning at 09:50 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
Cleaned enclosure AC filter.
2018-07-02 | Changed out BAM filter tape beginning at 10:15 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As left leak check = 0.5.
2018-07-05 | Cleaned BAM PM10 head.
Conducted site checks. OK.
2018-07-10 | Performed QC checks beginning at 09:15 CST.
Conducted site checks. OK.
2018-07-13 | Conducted site checks. OK.
2018-07-16 | Conducted site checks. OK.
2018-07-19 | Conducted site checks. OK.
Moved sample pump away from the data logger to reduce shelter temperature.
2018-07-24 | Conducted site checks. OK.
Checked with D. Szatkowski regarding the recent elevated shelter temperature.
Dave indicated the temperature is now within an allowable range.
2018-07-27 | Performed QC checks beginning at 10:15 CST.
Changed time from 10:17:05 to 10:16 CST.
Conducted site checks. OK.
2018-07-30 | Conducted site checks. OK.
2018-08-03 | Conducted site checks. OK.
2018-08-07 | Conducted site checks. OK.
2018-08-10 | Performed QC checks beginning at 09:40 CST.
Conducted site checks. OK.
2018-08-14 | Conducted site checks. OK.
Cleaned enclosure AC filter.
2018-08-17 | Found two small tears along the edge the the filter tape, similar those observed in Feb.
Advanced the tape, so the tears are well onto the take-up reel. Hour 10:00 CST was lost.
Conducted site checks. OK.
2018-08-23 | Performed QC checks beginning at 10:15 CST.
Conducted site checks. OK.
2018-08-28 | Conducted site checks. OK.
2018-08-30 | RAQIS reported a filter tape error at 05:00 CST. Found the tape broken with tears along
edges, like those on 8/17/18. Replaced the tape with a new one. This solution worked
in Feb. Hours 06:00 to 09:00 CST were lost.
2018-09-04 | Conducted site checks. OK.
2018-09-07 | Conducted site checks. OK.
2018-09-11 | Conducted site checks. OK.
2018-09-14 | Performed QC checks beginning at 09:35 CST.
Conducted site checks. OK.
2018-09-18 | Conducted site checks. OK.
2018-09-21 | Cleaned PM10 sampler head.
Conducted site checks. OK.
Changed time from 09:34:05 to 09:35 CST.
2018-09-24 | Performed QC checks beginning at 10:35 CST.
Conducted site checks. OK.
2018-09-27 | J. Sorkness performed a quarterly audit beginning at 10:50 CST, except for PM10 flow.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2018-09-28 | Performed PM10 audit flow checks with DeltaCal s/n 1301 beginning at 09:20 CST.
Results indicate the BAM flows to be within project limits.
2018-10-02 | Conducted site checks. OK.
2018-10-05 | Conducted site checks. OK.
2018-10-12 | Performed QC checks beginning at 09:15 CST.
Conducted site checks. OK.
2018-10-16 | Conducted site checks. OK.
2018-10-19 | Conducted site checks. OK.
2018-10-23 | Conducted site checks. OK.
2018-10-26 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2018-10-30 | Conducted site checks. OK.
Cleaned enclosure AC filter.
2018-11-02 | Changed out BAM filter tape beginning at 08:45 CST.
Leak check as found = 0.7, as left = 0.5.
Conducted site checks. OK.
2018-11-06 | Conducted site checks. OK.
2018-11-12 | Conducted site checks. OK.
2018-11-16 | Performed QC checks beginning at 10:35 CST.
Conducted site checks. OK.
2018-11-27 | Conducted site checks. OK.
2018-11-30 | Performed QC checks beginning at 11:15 CST.
Conducted site checks. OK.
2018-12-03 | Conducted site checks. OK.
2018-12-06 | J. Sorkness performed a quarterly audit beginning at 10:10 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2018-12-11 | Conducted site checks. OK.
2018-12-14 | Performed QC checks beginning at 11:10 CST.
Conducted site checks. OK.
2018-12-18 | Conducted site checks. OK.
2018-12-21 | Conducted site checks. OK.
Cleaned PM10 head.
2018-12-28 | Performed QC checks beginning at 10:20 CST.
Conducted site checks. OK.
2018-12-31 | Conducted site checks. OK.
2019-01-08 | Conducted site checks. OK.
Changed time from 11:29:05 to 11:28 CST.
2019-01-11 | Performed QC checks beginning at 10:25 CST.
Changed out BAM filter tape.
Leak check as found = 0.5, as left = 0.5.
Conducted site checks. OK.
2019-01-15 | Conducted site checks. OK.
2019-01-18 | Conducted site checks. OK.
2019-01-24 | Conducted site checks. OK.
2019-02-01 | Conducted site checks. OK.
2019-02-07 | Performed QC checks beginning at 13:40 CST.
Conducted site checks. OK.
2019-02-11 | Conducted site checks. OK.
2019-02-15 | Conducted site checks. OK.
2019-02-21 | Performed QC checks beginning at 14:20 CST.
Conducted site checks. OK.
Changed time from 14:20:05 to 14:19 CST.
2019-02-25 | Conducted site checks. OK.
2019-03-01 | Conducted site checks. OK.
2019-03-06 | Performed QC checks beginning at 12:40 CST.
Conducted site checks. OK.
2019-03-12 | Conducted site checks. OK.
2019-03-15 | Trimmed the end of the sample pump line at 09:40 CST.
Conducted site checks. OK.
2019-03-18 | Performed a leak check at 10:20 CST.
As found leak check = 0.5.
Conducted site checks. OK.
2019-03-20 | M. Wolfskill performed a quarterly audit beginning at 09:30 CST.
Results indicate all systems to be operating within allowable project limits.
Changed out BAM filter tape.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As found leak check = 0.5, As left leak check = 0.5.
Conducted site checks. OK.
2019-03-25 | Performed QC checks beginning at 10:10 CST.
Conducted site checks. OK.
2019-03-29 | Conducted site checks. OK.
2019-04-02 | Conducted site checks. OK.
Cleaned PM10 head.
2019-04-05 | Conducted site checks. OK.
2019-04-09 | Conducted site checks. OK.
2019-04-12 | Performed QC checks beginning at 10:15 CST.
Conducted site checks. OK.
2019-04-16 | Conducted site checks. OK.
Inspected enclosure HVAC filter. Found it clean.
2019-04-18 | Conducted site checks. OK.
2019-04-23 | Conducted site checks. OK.
2019-04-26 | Changed time from 09:51:05 to 09:50 CST.
Checked the DeltaCal flow.
It was reading 11.5, instead of the expected 15.0.
Conducted site checks. OK.
2019-04-30 | Performed QC checks beginning at 09:25 CST.
Conducted site checks. OK.
2019-05-03 | Conducted site checks. OK.
2019-05-06 | UPS battery backup failed earlier in the day. Replaced with a surge protector until
a new UPS arrives.
Conducted site checks. OK.
2019-05-10 | Conducted site checks. OK.
2019-05-16 | Performed QC checks beginning at 10:45 CST.
Conducted site checks. OK.
2019-05-22 | Conducted site checks. OK.
2019-05-24 | Conducted site checks. OK.
Changed out BAM filter tape at 10:35 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As found leak check = 0.5, As left leak check = 0.5.
Conducted site checks. OK.
2019-05-28 | Performed QC checks beginning at 11:30 CST.
Conducted site checks. OK.
2019-06-04 | Cleaned PM10 sampler head.
Conducted site checks. OK.
2019-06-07 | Power failure during hour 14:00 CST.
2019-06-11 | Performed QC checks beginning at 10:10 CST.
Conducted site checks. OK.
2019-06-14 | Conducted site checks. OK.
2019-06-18 | Cleaned shelter HVAC filter.
Conducted site checks. OK.
2019-06-21 | Conducted site checks. OK.
2019-06-26 | M. Wolfskill performed a quarterly audit beginning at 09:30 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2019-06-28 | Performed QC checks beginning at 09:15 CST.
Conducted site checks. OK.
2019-07-02 | Conducted site checks. OK.
2019-07-05 | Cleaned shelter HVAC filter.
Conducted site checks. OK.
2019-07-09 | Conducted site checks. OK.
2019-07-12 | Performed QC checks beginning at 09:15 CST.
Conducted site checks. OK.
2019-07-16 | Conducted site checks. OK.
2019-07-18 | Hour 14:00 has been invalidated by the BAM. D. Szatkowski checked and found a "D" flag
for that hour, which indicates a dirty cal membrane. I used compressed air to clean out the
area during a tape change on 7/26. Observed several tape fragments being blown out.
The m value went from 0.807 on 7/19 to 0.794 on 7/26, after the cleaning.
2019-07-19 | Conducted site checks. OK.
2019-07-23 | Cleaned PM10 sampler head.
Conducted site checks. OK.
2019-07-26 | Performed QC checks beginning at 10:20 CST.
Changed out BAM filter tape.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As found leak check = 0.5, As left leak check = 0.5.
Conducted site checks. OK.
2019-07-30 | Conducted site checks. OK.
2019-08-02 | Conducted site checks. OK.
2019-08-06 | Conducted site checks. OK.
2019-08-09 | Performed QC checks beginning at 10:20 CST.
Conducted site checks. OK.
2019-08-13 | Conducted site checks. OK.
2019-08-16 | Conducted site checks. OK.
2019-08-20 | Cleaned shelter HVAC filter.
Changed time from 09:48:05 to 09:47 CST.
Conducted site checks. OK.
2019-08-23 | Performed QC checks beginning at 10:15 CST.
Conducted site checks. OK.
2019-08-24 | A filter tape error was recorded for hour 09:00. Likely due to a machine malfunction. I
inspected the tape system on 8/27 and everything appeared normal.
2019-08-27 | Conducted site checks. OK.
2019-09-03 | Conducted site checks. OK.
2019-09-05 | We have been having intermittent com issues. B. Stringer is sending a replacement router.
Conducted site checks. OK.
2019-09-06 | Installed the router. B. Stringer indicated the coms are working well.
2019-09-10 | Performed QC checks beginning at 10:15 CST.
Conducted site checks. OK.
2019-09-11 | Hours 03:00 and 04:00 CST are invalid due to a power failure.
2019-09-13 | We are still having occasional com issues. B. Stringer is sending a replacement
Lantronix.
Conducted site checks. OK.
2019-09-17 | Installed the new Lantronix. Power cycled the BAM at 10:10 CST during installation.
That hour is lost.
2019-09-19 | M. Wolfskill performed a quarterly audit beginning at 08:20 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2019-09-24 | Conducted site checks. OK.
2019-09-27 | Performed QC checks beginning at 10:20 CST.
Conducted site checks. OK.
2019-09-30 | Changed out BAM filter tape.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As found leak check = 0.5, As left leak check = 0.5.
Conducted site checks. OK.
2019-10-04 | Conducted site checks. OK.
2019-10-07 | Conducted site checks. OK.
2019-10-09 | Conducted site checks. OK.
2019-10-15 | Performed QC checks beginning at 10:20 CST.
Changed time from 10:16:05 to 10:15 CST.
Conducted site checks. OK.
2019-10-21 | Conducted site checks. OK.
2019-10-24 | Conducted site checks. OK.
2019-10-28 | Performed QC checks beginning at 12:20 CST.
Conducted site checks. OK.

## ST-NE
2014-10-07 | Performed BAM QC checks starting at 11:40 CST.
Cleaned PM10 sample head and nozzle.
Leak check as found was 0.5, after nozzle cleaning it was 0.5.
Conducted site checks. OK.
Checked temperature inside ST-NE enclosure. Data logger = 27.1, Delta Cal = 24.7 deg C.
Reset the onsite PC. It would not connect to the Partisols and the icons had double
vision.
2014-10-10 | Conducted site checks. OK.
2014-10-13 | Conducted site checks. OK.
2014-10-16 | Conducted site checks. OK.
2014-10-21 | Performed BAM QC checks starting at 11:00 CST.
Conducted site checks. OK.
2014-10-24 | Conducted site checks. OK.
Took BAM offline to change time from 09:43:20 to 09:42.
2014-10-28 | Conducted site checks. OK.
2014-10-30 | Conducted site checks. OK.
2014-11-03 | Performed BAM QC checks starting at 13:45 CST.
Conducted site checks. OK.
2014-11-05 | Conducted site checks. OK.
2014-11-10 | Conducted site checks. OK.
2014-11-14 | Conducted site checks. OK.
Shelter temperature reported as too high. Adjusted thermostat from 80 to 70 degrees F.
2014-11-17 | Conducted site checks. OK.
2014-11-21 | Performed BAM QC checks starting at 12:15 CST.
Replaced filter tape.
Conducted site checks. OK.
2014-11-26 | Conducted site checks. OK.
2014-12-01 | Conducted site checks. OK.
2014-12-05 | Performed BAM QC checks starting at 12:25 CST.
Conducted site checks. OK.
Cleaned PM10 sample head.
2014-12-08 | J. Sorkness conducted audits of the BAM.
Conducted site checks. OK.
2014-12-12 | Conducted site checks. OK.
B. Stringer uploaded a new operating system to the spare data logger that had been
powered by the existing data logger at the ST-NE site since November 26. This connection
caused the voltages recorded on the ST-NE data logger to fluctuate. In the future,
any additional data loggers connected to the site laptop for purpose of updating the
operating system software will be powered by its own power supply.
The spare data logger has been removed from the ST-NE site. As a result, data recorded
from 11/26 to 12/12 on the ST-NE data logger was affected. PM10 data recorded by the BAM
should be unaffected and are available as a substitute for the affected data.
2014-12-16 | Conducted site checks. OK.
Changed time from 10:57:00 to 10:56 CST.
2014-12-18 | Performed BAM QC checks starting at 10:00 CST.
Conducted site checks. OK.
Attached the spare data logger to the site laptop so the operating system can be updated.
This data logger has its own power supply.
2014-12-23 | Conducted site checks. OK.
2014-12-30 | Conducted site checks. OK.
2015-01-02 | Conducted site checks. OK.
2015-01-05 | Conducted site checks. OK.
2015-01-08 | Conducted site checks. OK.
2015-01-12 | Perform QC checks beginning at 12:20 CST.
Conducted site checks. OK.
2015-01-13 | Begin annual zeroing at 10:25 CST.
Conducted site checks. OK.
Changed units' background concentration to 0.0000.
2015-01-16 | Conducted site checks. OK.
Concluded annual zeroing at 13:05 CST.
Changed units' background concentration to -0.0033.
2015-01-20 | Input new zero background concentration of -0.0035 into BAM at 11:05 CST.
Conducted site checks. OK.
2015-01-23 | Conducted site checks. OK.
2015-01-26 | Perform QC checks beginning at 11:30 CST.
Conducted site checks. OK.
2015-01-29 | Conducted site checks. OK.
2015-01-31 | Site experienced a power failure at 07:00 CST. Power was restored at 09:50 CST.
As a result, hours 07:00, 08:00, and 09:00 are invalid.
2015-02-04 | Conducted site checks. OK.
2015-02-06 | Conducted site checks. OK.
Changed time from 10:55:05 to 10:54 CST.
2015-02-10 | Perform QC checks beginning at 13:10 CST.
Conducted site checks. OK.
2015-02-13 | Conducted site checks. OK.
Todd Chandler mentioned that the conduit to the shelter was tested by an electrician
and found to have a fluctuating current. This may explain some of the invalid hours we
have been experiencing the last month or so. KCBX plans to replace the existing conduit
and wiring with a new one. Installation is planned to begin next week and last one to
one and one half days. During this change out, an electrical generator will be operated to
maintain power to the site. I mentioned that the generator should be located downwind
and as far away as possible from the monitor.
2015-02-16 | Conducted site checks. OK.
Rewiring of site delayed due to missing parts that did not come in.
2015-02-19 | Conducted site checks. OK.
2015-02-24 | Conducted site checks. OK.
2015-02-27 | Perform QC checks beginning at 13:05 CST.
Conducted site checks. OK.
2015-03-02 | Conducted site checks. OK.
2015-03-06 | Changed filter tape beginning at 14:30 CST.
Performed leak check = 0.3.
Cleaned PM10 head.
Conducted site checks. OK.
2015-03-09 | Conducted site checks. OK.
2015-03-11 | Conducted site checks. OK.
2015-03-17 | Conducted site checks. OK.
2015-03-19 | J. Sorkness performed audits beginning at 12:40 CST.
2015-03-20 | Conducted site checks. OK.
2015-03-24 | Perform QC checks beginning at 11:20 CST.
Conducted site checks. OK.
2015-03-27 | Conducted site checks. OK.
2015-03-31 | Perform leak checks beginning at 12:20 CST.
Leak as found = 0.6, as left = 0.5.
Conducted site checks. OK.
2015-04-01 | Conducted site checks. OK.
2015-04-08 | Perform QC checks beginning at 12:30 CST.
Conducted site checks. OK.
2015-04-10 | Conducted site checks. OK.
2015-04-15 | Conducted site checks. OK.
2015-04-17 | Conducted site checks. OK.
2015-04-24 | Perform QC checks beginning at 11:10 CST.
Conducted site checks. OK.
2015-04-27 | Conducted site checks. OK.
Installed new alarm wiring beginning at 12:30 CST.
2015-04-30 | Conducted site checks. OK.
Changed filter tape beginning at 14:10 CST.
2015-05-04 | Perform QC checks beginning at 11:05 CST.
Conducted site checks. OK.
2015-05-07 | Conducted site checks. OK.
2015-05-11 | Conducted site checks. OK.
2015-05-14 | Conducted site checks. OK.
2015-05-18 | Conducted site checks. OK.
2015-05-21 | Perform QC checks beginning at 11:50 CST.
Conducted site checks. OK.
2015-05-27 | Conducted site checks. OK.
2015-05-29 | Conducted site checks. OK.
2015-06-01 | Conducted site checks. OK.
2015-06-04 | Perform QC checks beginning at 13:00 CST.
Changed time from 13:07:05 to 13:06 CST.
Cleaned nozzle and screen.
Cleaned PM10 head.
Conducted site checks. OK.
2015-06-09 | Conducted site checks. OK.
2015-06-12 | Conducted site checks. OK.
2015-06-15 | Performed leak check beginning at 12:40 CST.
Leak check = 0.5.
Conducted site checks. OK.
2015-06-17 | J. Sorkness performed audit checks beginning at 08:30 CST.
2015-06-18 | Conducted site checks. OK.
2015-06-22 | Conducted site checks. OK.
2015-06-24 | Conducted site checks. OK.
Changed filter tape at 08:20 CST.
2015-06-30 | Conducted site checks. OK.
2015-07-02 | Performed QC checks beginning at 07:40 CST.
Conducted site checks. OK.
2015-07-07 | Conducted site checks. OK.
2015-07-10 | Conducted site checks. OK.
2015-07-14 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
2015-07-16 | Conducted site checks. OK.
2015-07-22 | Conducted site checks. OK.
2015-07-24 | Conducted site checks. OK.
2015-07-28 | Performed QC checks beginning at 11:25 CST.
Conducted site checks. OK.
2015-07-30 | Conducted site checks. OK.
2015-08-05 | Conducted site checks. OK.
Changed time from 10:49:20 to 10:48 CST.
2015-08-07 | Conducted site checks. OK.
2015-08-11 | Performed a BAM leak check beginning at 11:40 CST.
Conducted site checks. OK.
2015-08-12 | J. Sorkness performed a quarterly audit beginning at 13:10 CST.
Conducted site checks. OK.
2015-08-14 | Conducted site checks. OK.
2015-08-19 | Conducted site checks. OK.
2015-08-21 | Performed QC checks beginning at 11:15 CST.
Changed BAM filter tape.
Conducted site checks. OK.
2015-08-28 | Conducted site checks. OK.
Took BAM offline to record Last M and ABS readings at 10:40 CST.
Last M = 0.798, ABS = 0.789
2015-09-01 | Conducted site checks. OK.
2015-09-04 | Performed QC checks beginning at 09:55 CST.
Conducted site checks. OK.
2015-09-09 | Conducted site checks. OK.
2015-09-11 | Conducted site checks. OK.
Cleaned PM10 head.
2015-09-15 | Conducted site checks. OK.
2015-09-16 | Replaced the sample pump with a rebuilt one. Hour 13:00 CST is invalid due to the change
out.
Conducted site checks. OK.
2015-09-22 | Performed QC checks beginning at 10:35 CST.
Began 72-hour zero calibration following QC checks.
Changed zero offset to 0.0000.
Conducted site checks. OK.
2015-09-25 | Concluded zero calibration.
Adjusted zero offset to former number of -0.0035.
New zero offset numbers will be input next week.
Valid data collection resumed at 12:00 CST.
Conducted site checks. OK.
2015-09-28 | Placed new zero offset into monitor at 11:25 CST.
The new zero offset is -0.0017.
Conducted site checks. OK.
2015-09-30 | Conducted site checks. OK.
2015-10-02 | Conducted site checks. OK.
2015-10-05 | Performed QC checks beginning at 12:50 CST.
Conducted site checks. OK.
Changed time from 12:56:10 to 12:55 CST.
Brandon contacted me about an electrical circuit change at the ST-NE site.
A KCBX electrician changed the circuit to the ComEd emergency service system. The
electrician indicated the new service will be more reliable. Data from 13:40 to 14:15 are
invalid due to the power failure during the change over.
2015-10-08 | Conducted site checks. OK.
2015-10-13 | Conducted site checks. OK.
2015-10-15 | Conducted site checks. OK.
2015-10-19 | Conducted site checks. OK.
2015-10-22 | Performed QC checks beginning at 13:50 CST.
Changed filter tape.
Conducted site checks. OK.
2015-10-27 | Conducted site checks. OK.
2015-10-29 | Conducted site checks. OK.
2015-11-02 | Conducted site checks. OK.
2015-11-04 | Performed QC checks beginning at 15:30 CST.
Conducted site checks. OK.
2015-11-10 | Conducted site checks. OK.
2015-11-12 | Conducted site checks. OK.
2015-11-16 | Performed a leak check beginning at 13:20 CST.
Leak = 0.5.
Conducted site checks. OK.
2015-11-18 | J. Sorkness performed a quarterly audit beginning at 11:20 CST.
Conducted site checks. OK.
2015-11-23 | Conducted site checks. OK.
2015-11-25 | Conducted site checks. OK.
2015-11-27 | Unidentified 985 alarms generated for hours 10:00, 11:00, and 17:00 CST.
2015-11-30 | Performed QC checks beginning at 13:30 CST.
Found display has vertical lines. Harder to read. Jim has ordered a replacement.
Conducted site checks. OK.
2015-12-03 | Conducted site checks. OK.
2015-12-08 | Conducted site checks. OK.
2015-12-10 | Conducted site checks. OK.
Cleaned PM10 head.
2015-12-14 | Conducted site checks. OK.
2015-12-16 | Performed QC checks beginning at 11:35 CST.
Replaced the BAM front door with the display.
Changed time from 13:18:20 to 13:17 CST.
Found water inside the BAM door. This may have caused the display to fail.
The front of the enclosure has sunken a bit. This is causing water to pool near the front
of the enclosure following rain, possibly giving it a longer time to seep inside.
During my next visit, I will raise the front of the enclosure and block it, so the water runs
to the back. In addition, the seals on top will be recaulked.
Conducted site checks. OK.
2015-12-22 | Changed filter tape beginning at 12:10 CST.
Conducted site checks. OK.
2015-12-29 | Performed QC checks beginning at 11:30 CST.
Conducted site checks. OK.
2015-12-31 | Conducted site checks. OK.
Found the display to be hard to read. Found a small amount of water inside the enclosure.
Raised the front of the enclosure, so water will run off the back. We have had several days
of continuous rain. Placed paper towels on the top of the monitor to soak up any water.
Will replace the display as soon as possible.
2016-01-04 | Conducted site checks. OK.
Replaced BAM display. Recaulked flange surrounding PM10 inlet.
2016-01-06 | Conducted site checks. OK.
2016-01-11 | Conducted site checks. OK.
2016-01-14 | Conducted site checks. OK.
2016-01-19 | Performed QC checks beginning at 11:45 CST.
Conducted site checks. OK.
2016-01-20 | Power failure from 17:00 to 19:00 CST.
2016-01-21 | Conducted site checks. OK.
2016-01-25 | Conducted site checks. OK.
2016-01-27 | Conducted site checks. OK.
2016-02-01 | Conducted site checks. OK.
2016-02-04 | Conducted site checks. OK.
2016-02-08 | Performed QC checks beginning at 11:05 CST.
Changed filter tape.
Conducted site checks. OK.
2016-02-11 | Conducted site checks. OK.
2016-02-16 | Conducted site checks. OK.
Changed time from 11:50:05 to 11:49 CST.
2016-02-19 | Conducted site checks. OK.
2016-02-23 | Performed QC checks beginning at 11:20 CST.
Conducted site checks. OK.
2016-02-26 | Conducted site checks. OK.
2016-02-29 | Conducted site checks. OK.
2016-03-04 | Conducted site checks. OK.
2016-03-08 | Performed QC checks beginning at 10:50 CST.
Conducted site checks. OK.
2016-03-11 | Conducted site checks. OK.
Cleaned PM10 head.
2016-03-14 | Conducted site checks. OK.
2016-03-17 | Conducted site checks. OK.
Cleaned enclosure air conditioner and changed filter.
2016-03-23 | J. Sorkness performed a quarterly audit beginning at 11:20 CST.
Conducted site checks. OK.
2016-03-25 | Conducted site checks. OK.
2016-04-01 | Conducted site checks. OK.
2016-04-04 | Performed QC checks beginning at 11:10 CST.
Conducted site checks. OK.
2016-04-07 | Conducted site checks. OK.
Changed time from 11:22:05 to 11:21 CST.
2016-04-11 | Performed preventative maintenace beginning at 11:45 CST.
Cleaned nozzle, screen, and tape rollers.
Changed filter tape.
Conducted site checks. OK.
2016-04-14 | Conducted site checks. OK.
2016-04-18 | Conducted site checks. OK.
Visible fugitive dust being generated by trucks on Burley Avenue.
2016-04-21 | Performed QC checks beginning at 11:35 CST.
Conducted site checks. OK.
2016-04-25 | Conducted site checks. OK.
2016-04-29 | Conducted site checks. OK.
Cleaned PM10 head.
2016-05-03 | Performed QC checks beginning at 11:10 CST.
Conducted site checks. OK.
Last M = 0.792, ABS = 0.789.
2016-05-05 | Conducted site checks. OK.
2016-05-09 | Conducted site checks. OK.
2016-05-12 | Conducted site checks. OK.
2016-05-16 | Performed QC checks beginning at 11:40 CST.
Conducted site checks. OK.
2016-05-19 | Conducted site checks. OK.
2016-05-24 | Conducted site checks. OK.
2016-05-25 | Elevated concentrations recorded due to high winds. Hour 16:00 experienced high loading.
We believe values reading 986 are true readings representing a concentration of 986 or higher.
2016-05-27 | Performed QC checks beginning at 09:45 CST.
Conducted site checks. OK.
2016-05-28 | Power failure during hour 21:00.
2016-05-31 | Conducted site checks. OK.
2016-06-02 | Conducted site checks. OK.
Cleaned BAM enclosure air conditioner filter.
2016-06-07 | J. Sorkness performed a quarterly audit beginning at 11:00 CST.
Conducted site checks. OK.
2016-06-10 | Conducted site checks. OK.
Changed out BAM filter tape at 07:15 CST,
2016-06-14 | Performed QC checks beginning at 14:35 CST.
Conducted site checks. OK.
2016-06-20 | Cleaned PM10 head.
Conducted site checks. OK.
2016-06-23 | Cleaned enclosure air conditioning filter.
Conducted site checks. OK.
Observed a contractor spraying weed killer near the site.
2016-06-27 | Performed QC checks beginning at 09:45 CST.
Conducted site checks. OK.
2016-06-30 | Conducted site checks. OK.
2016-07-06 | Conducted site checks. OK.
2016-07-08 | Cleaned PM10 head.
Conducted site checks. OK.
2016-07-11 | Performed QC checks beginning at 11:25 CST.
Conducted site checks. OK.
2016-07-14 | Conducted site checks. OK.
2016-07-19 | Conducted site checks. OK.
2016-07-22 | Conducted site checks. OK.
2016-07-26 | Performed QC checks beginning at 10:15 CST.
Conducted site checks. OK.
2016-07-28 | Changed out filter tape beginning at 12:30 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
Conducted site checks. OK.
2016-08-02 | Conducted site checks. OK.
2016-08-05 | Conducted site checks. OK.
Changed time from 10:22:05 to 10:21 CST.
2016-08-09 | Performed QC checks beginning at 10:20 CST.
Conducted site checks. OK.
2016-08-12 | Conducted site checks. OK.
2016-08-15 | Conducted site checks. OK.
2016-08-18 | Conducted site checks. OK.
2016-08-22 | Performed QC checks beginning at 10:45 CST.
Conducted site checks. OK.
2016-08-26 | Conducted site checks. OK.
Cleaned shelter air condtioner filter.
2016-08-29 | Conducted site checks. OK.
2016-09-01 | J. Sorkness performed a quarterly audit beginning at 10:15 CST.
Conducted site checks. OK.
2016-09-06 | Conducted site checks. OK.
2016-09-09 | Conducted site checks. OK.
Cleaned PM10 head.
2016-09-13 | Performed QC checks beginning at 10:40 CST.
Conducted site checks. OK.
2016-09-16 | Conducted site checks. OK.
2016-09-19 | Conducted site checks. OK.
2016-09-22 | Conducted site checks. OK.
Changed out filter tape starting at 08:50 CST.
2016-10-01 | Conducted site checks. OK.
Changed time from 10:03:15 to 10:02 CST.
2016-10-04 | Conducted site checks. OK.
2016-10-07 | Conducted site checks. OK.
2016-10-11 | Performed QC checks beginning at 11:35 CST.
Conducted site checks. OK.
2016-10-14 | Conducted site checks. OK.
Cleaned PM10 sample head.
2016-10-17 | Conducted site checks. OK.
2016-10-20 | Conducted site checks. OK.
2016-10-25 | Performed QC checks beginning at 10:35 CST.
Conducted site checks. OK.
Last M = 0.794, ABS = 0.789.
2016-10-28 | Conducted site checks. OK.
2016-11-02 | Conducted site checks. OK.
2016-11-04 | Conducted site checks. OK.
Cleaned enclosure AC filter.
2016-11-08 | Conducted site checks. OK.
2016-11-11 | Performed QC checks beginning at 10:50 CST.
Conducted site checks. OK.
2016-11-15 | Conducted site checks. OK.
2016-11-18 | Conducted site checks. OK.
2016-11-22 | Performed QC checks beginning at 11:45 CST.
Changed out filter tape.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
Conducted site checks. OK.
2016-11-25 | Conducted site checks. OK.
2016-11-30 | J. Sorkness performed a quarterly audit beginning at 15:35 CST.
Results indicate all systems to be operating within allowable project limits.
Time changed from 14:21:05 to 14:20 CST.
Conducted site checks. OK.
2016-12-02 | Conducted site checks. OK.
2016-12-06 | Conducted site checks. OK.
2016-12-09 | Conducted site checks. OK.
2016-12-13 | Performed QC checks beginning at 11:20 CST.
Conducted site checks. OK.
2016-12-16 | Conducted site checks. OK.
2016-12-20 | Conducted site checks. OK.
2016-12-22 | Conducted site checks. OK.
Cleaned PM10 head.
2016-12-28 | Performed QC checks beginning at 13:45 CST.
Conducted site checks. OK.
2016-12-30 | Conducted site checks. OK.
2017-01-03 | Conducted site checks. OK.
Flow error from 15:00-16:00 CST.
2017-01-06 | Conducted site checks. OK.
2017-01-09 | Conducted site checks. OK.
2017-01-13 | Performed QC checks beginning at 11:25 CST.
Conducted site checks. OK.
2017-01-17 | Flow error at 02:00 CST.
Performed filter tape change out beginning at 11:40 CST.
Conducted site checks. OK.
2017-01-20 | Flow error at 10:00 CST.
2017-01-23 | Conducted site checks. OK.
Changed time from 10:52:10 to 10:51 CST.
Occasional flow errors have been recorded on this unit. Inspected the vacuum pump and
tubing. No obvious signs of damage or wear. Cut off the end of the tubing inserted into
the pump and reattached. This end is exposed to high temperatures from the pump and
may have developed a slight leak.
2017-01-27 | Performed QC checks beginning at 11:30 CST.
Conducted site checks. OK.
2017-01-31 | Conducted site checks. OK.
2017-02-06 | Conducted site checks. OK.
2017-02-13 | Performed QC checks beginning at 11:30 CST.
Conducted site checks. OK.
2017-02-17 | Conducted site checks. OK.
Power failure at 09:45 CST due to the KCBX electrician disconnecting the Partisol
samplers from the circuit breaker.
2017-02-23 | Conducted site checks. OK.
2017-02-27 | Performed QC checks beginning at 10:25 CST.
Conducted site checks. OK.
2017-03-03 | Conducted site checks. OK.
2017-03-08 | Cleaned PM10 heads.
Conducted site checks. OK.
2017-03-16 | Performed QC checks beginning at 09:25 CST.
Conducted site checks. OK.
2017-03-20 | Cleaned enclosure AC filters.
Conducted site checks. OK.
2017-03-24 | J. Sorkness performed a quarterly audit beginning at 09:50 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2017-03-27 | Changed out BAM filter tape beginning at 09:30 CST.
Conducted site checks. OK.
2017-03-31 | Performed QC checks beginning at 09:45 CST.
Conducted site checks. OK.
2017-04-01 | Power failure from 09:00 to 10:00 CST.
2017-04-07 | Conducted site checks. OK.
2017-04-10 | Conducted site checks. OK.
2017-04-13 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2017-04-18 | Conducted site checks. OK.
2017-04-21 | Conducted site checks. OK.
2017-04-25 | Performed QC checks beginning at 10:10 CST.
Conducted site checks. OK.
2017-04-28 | Conducted site checks. OK.
2017-05-05 | Conducted site checks. OK.
2017-05-12 | Performed QC checks beginning at 09:20 CST.
Changed time from 09:23:10 to 09:22 CST.
Conducted site checks. OK.
2017-05-16 | Conducted site checks. OK.
2017-05-19 | Conducted site checks. OK.
Cleaned PM10 heads.
2017-05-23 | Conducted site checks. OK.
2017-05-26 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
Last M = 0.795, ABS = 0.789.
2017-05-30 | Changed out filter tape beginning at 10:10 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
Leak check as found = 0.4. Leak check as left = 0.5
Conducted site checks. OK.
2017-06-05 | Conducted site checks. OK.
2017-06-08 | Performed QC checks beginning at 10:10 CST.
Conducted site checks. OK.
2017-06-13 | Conducted site checks. OK.
2017-06-20 | Conducted site checks. OK.
2017-06-23 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
Cleaned shelter AC filter.
2017-06-29 | J. Sorkness performed a quarterly audit beginning at 09:50 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2017-07-05 | Conducted site checks. OK.
2017-07-11 | Performed QC checks beginning at 10:30 CST.
Conducted site checks. OK.
2017-07-14 | Conducted site checks. OK.
Changed time from 09:50:05 to 09:49 CST.
2017-07-18 | Conducted site checks. OK.
Cleaned PM10 head.
2017-07-21 | Conducted site checks. OK.
2017-07-24 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2017-07-27 | Conducted site checks. OK.
2017-07-31 | Conducted site checks. OK.
2017-08-04 | Conducted site checks. OK.
Changed out BAM filter tape at 12:20 CST.
2017-08-08 | Performed QC checks beginning at 09:35 CST.
Conducted site checks. OK.
2017-08-11 | Conducted site checks. OK.
2017-08-15 | Conducted site checks. OK.
2017-08-16 | Conducted site checks. OK.
Replaced shelter AC filter.
2017-08-22 | Conducted site checks. OK.
2017-08-25 | Performed QC checks beginning at 09:25 CST.
Conducted site checks. OK.
2017-08-29 | Conducted site checks. OK.
2017-08-31 | Conducted site checks. OK.
2017-09-05 | Conducted site checks. OK.
Cleaned PM10 head.
2017-09-08 | J. Sorkness performed a quarterly audit beginning at 07:35 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2017-09-12 | Performed QC checks beginning at 10:30 CST.
Conducted site checks. OK.
2017-09-15 | Conducted site checks. OK.
2017-09-19 | Conducted site checks. OK.
2017-09-22 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2017-10-02 | Conducted site checks. OK.
2017-10-05 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
2017-10-09 | Changed out BAM filter tape at 09:30.
Conducted site checks. OK.
2017-10-12 | Conducted site checks. OK.
2017-10-16 | Conducted site checks. OK.
2017-10-20 | Conducted site checks. OK.
2017-10-24 | Conducted site checks. OK.
2017-10-27 | Performed QC checks beginning at 10:20 CST.
Conducted site checks. OK.
2017-11-06 | Conducted site checks. OK.
Changed time from 11:41:05 to 11:40 CST.
2017-11-10 | Conducted site checks. OK.
2017-11-14 | Rebuilt the BAM sample pump. Offline at 11:50 CST for the change out.
Conducted site checks. OK.
2017-11-17 | Performed QC checks beginning at 08:55 CST.
Conducted site checks. OK.
2017-11-24 | Conducted site checks. OK.
2017-11-27 | Performed QC checks beginning at 11:30 CST.
Conducted site checks. OK.
2017-11-30 | J. Sorkness performed a quarterly audit beginning at 08:10 CST.
Results indicate all systems to be operating within allowable project limits.
Adjusted temperature from 5.5 to 6.7 degrees C.
Conducted site checks. OK.
2017-12-05 | Conducted site checks. OK.
2017-12-12 | Conducted site checks. OK.
2017-12-15 | Performed QC checks beginning at 11:30 CST.
Conducted site checks. OK.
2017-12-19 | Changed out BAM filter tape beginning at 10:15 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
Leak check as found = 0.5, as left = 0.5.
Conducted site checks. OK.
2017-12-22 | Conducted site checks. OK.
2017-12-26 | Conducted site checks. OK.
2017-12-29 | Conducted site checks. OK.
2018-01-02 | The monitor recorded flow errors from 05:00 to 07:00 CST, likely due to extreme low
temperatures. The BAM resumed normal operation at 08:00 CST.
Concentrations were recorded by the BAM. D. Szatkowski will check with MetOne to
see if these concentrations are valid.
2018-01-03 | Performed QC checks beginning at 11:10 CST.
Conducted site checks. OK.
Changed time from 11:09:10 to 11:08 CST.
2018-01-09 | Conducted site checks. OK.
2018-01-12 | Conducted site checks. OK.
2018-01-17 | Conducted site checks. OK.
2018-01-19 | Performed QC checks beginning at 09:50 CST.
Conducted site checks. OK.
2018-01-22 | Conducted site checks. OK.
2018-01-25 | Conducted site checks. OK.
2018-01-29 | Performed QC checks beginning at 11:25 CST.
Conducted site checks. OK.
2018-02-02 | Conducted site checks. OK.
2018-02-06 | Conducted site checks. OK.
2018-02-13 | Conducted site checks. OK.
2018-02-16 | Conducted site checks. OK.
2018-02-20 | Performed QC checks beginning at 11:30 CST.
Conducted site checks. OK.
2018-02-23 | Changed out BAM filter tape beginning at 09:35 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
2018-02-26 | Conducted site checks. OK.
2018-03-01 | Conducted site checks. OK.
2018-03-06 | Performed QC checks beginning at 11:30 CST.
Changed time from 11:28:05 to 11:27 CST.
Conducted site checks. OK.
2018-03-09 | Conducted site checks. OK.
Changed time from 10:43:05 to 10:42 CST.
2018-03-13 | Conducted site checks. OK.
2018-03-16 | Conducted site checks. OK.
2018-03-20 | Performed QC checks beginning at 10:40 CST.
Conducted site checks. OK.
2018-03-23 | Conducted site checks. OK.
Changed time from 09:53:10 to 09:52 CST.
2018-03-27 | Conducted site checks. OK.
2018-03-29 | J. Sorkness performed a quarterly audit beginning at 10:10 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2018-04-03 | Conducted site checks. OK.
2018-04-06 | Conducted site checks. OK.
2018-04-10 | Performed QC checks beginning at 10:30 CST.
Conducted site checks. OK.
2018-04-13 | Conducted site checks. OK.
2018-04-16 | Conducted site checks. OK.
2018-04-19 | Conducted site checks. OK.
2018-04-23 | Performed QC checks beginning at 10:35 CST.
Conducted site checks. OK.
2018-04-26 | Changed out BAM filter tape beginning at 11:45 CST.
Leak check as found = 0.4, as left = 0.5.
Conducted site checks. OK.
2018-05-01 | Conducted site checks. OK.
2018-05-04 | Conducted site checks. OK.
2018-05-11 | Performed QC checks beginning at 13:20 CST.
Conducted site checks. OK.
2018-05-18 | Conducted site checks. OK.
2018-05-22 | Conducted site checks. OK.
Cleaned enclosure AC filters.
2018-05-25 | Performed QC checks beginning at 08:40 CST.
Conducted site checks. OK.
2018-05-29 | Conducted site checks. OK.
2018-06-01 | Conducted site checks. OK.
2018-06-05 | Conducted site checks. OK.
Cleaned enclosure AC filter.
2018-06-08 | Conducted site checks. OK.
2018-06-11 | Performed QC checks beginning at 13:40 CST.
Conducted site checks. OK.
2018-06-14 | Conducted site checks. OK.
2018-06-18 | Conducted site checks. OK.
2018-06-21 | Conducted site checks. OK.
2018-06-25 | Performed QC checks beginning at 09:25 CST.
Conducted site checks. OK.
2018-06-28 | J. Sorkness performed a quarterly audit beginning at 10:10 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
Cleaned enclosure AC filter.
2018-06-29 | Hour 07:00 CST is invalid due to a power failure caused by an HVAC technician inspecting
the unit.
2018-07-02 | Changed out BAM filter tape beginning at 09:30 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As left leak check = 0.4.
2018-07-05 | Cleaned BAM PM10 head.
Conducted site checks. OK.
2018-07-10 | Performed QC checks beginning at 09:25 CST.
Conducted site checks. OK.
2018-07-13 | Conducted site checks. OK.
2018-07-16 | Conducted site checks. OK.
2018-07-19 | Conducted site checks. OK.
2018-07-24 | Conducted site checks. OK.
2018-07-27 | Performed QC checks beginning at 10:30 CST.
Changed time from 10:30:05 to 10:29 CST.
Conducted site checks. OK.
2018-07-30 | Conducted site checks. OK.
2018-08-03 | Conducted site checks. OK.
2018-08-07 | Conducted site checks. OK.
2018-08-10 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2018-08-14 | Conducted site checks. OK.
Cleaned enclosure AC filter.
2018-08-17 | Conducted site checks. OK.
2018-08-23 | Performed QC checks beginning at 09:20 CST.
Conducted site checks. OK.
2018-08-28 | Conducted site checks. OK.
2018-08-30 | Conducted site checks. OK.
2018-09-04 | Conducted site checks. OK.
2018-09-07 | Changed out BAM filter tape beginning at 09:30 CST.
Leak check as found = 0.5, as left = 0.5.
Conducted site checks. OK.
2018-09-11 | Conducted site checks. OK.
2018-09-14 | Performed QC checks beginning at 09:25 CST.
Conducted site checks. OK.
2018-09-18 | Conducted site checks. OK.
2018-09-21 | Cleaned PM10 sampler head.
Conducted site checks. OK.
2018-09-24 | Performed QC checks beginning at 10:05 CST.
Conducted site checks. OK.
Changed time from 10:04:05 to 10:03 CST.
2018-09-27 | J. Sorkness performed a quarterly audit beginning at 09:50 CST, except for PM10 flow.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2018-09-28 | Performed PM10 audit flow checks with DeltaCal s/n 1301 beginning at 09:30 CST.
Results indicate the BAM flows to be within project limits.
2018-10-02 | Conducted site checks. OK.
2018-10-05 | Conducted site checks. OK.
2018-10-12 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2018-10-16 | Conducted site checks. OK.
2018-10-19 | Conducted site checks. OK.
2018-10-23 | Conducted site checks. OK.
2018-10-26 | Performed QC checks beginning at 08:20 CST.
Conducted site checks. OK.
2018-10-30 | Conducted site checks. OK.
Cleaned enclosure AC filter.
2018-11-02 | Conducted site checks. OK.
2018-11-06 | Conducted site checks. OK.
2018-11-12 | Changed out BAM filter tape beginning at 10:35 CST.
Leak check as found = 0.7, as left = 0.5.
Conducted site checks. OK.
2018-11-16 | Performed QC checks beginning at 10:25 CST.
Conducted site checks. OK.
2018-11-27 | Conducted site checks. OK.
2018-11-30 | Performed QC checks beginning at 11:35 CST.
Conducted site checks. OK.
2018-12-03 | Conducted site checks. OK.
2018-12-06 | J. Sorkness performed a quarterly audit beginning at 10:35 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2018-12-11 | Conducted site checks. OK.
2018-12-14 | Performed QC checks beginning at 10:35 CST.
Conducted site checks. OK.
2018-12-18 | Conducted site checks. OK.
2018-12-21 | Conducted site checks. OK.
Cleaned PM10 head.
2018-12-28 | Performed QC checks beginning at 10:30 CST.
Changed time from 10:31:05 to 10:30 CST.
Conducted site checks. OK.
2018-12-31 | Conducted site checks. OK.
2019-01-08 | Conducted site checks. OK.
2019-01-11 | Performed QC checks beginning at 10:10 CST.
Conducted site checks. OK.
2019-01-15 | Conducted site checks. OK.
2019-01-18 | Changed out BAM filter tape beginning at 10:10 CST.
Leak check as found = 0.6, as left = 0.5.
Conducted site checks. OK.
2019-01-24 | Conducted site checks. OK.
2019-01-30 | Filter tape error reported for hour 08:00 CST.
Investigated this issue on 2/1 and found the tape on the collection side reel to have a
slight slack in it. Talked to D. Szatkowski about the problem. Advanced the tape
several feet and moved the tension roller back and forth several times in case it is sticking.
Performed a self test and the tension seems normal again.
2019-02-01 | Conducted site checks. OK.
Hour 11:00 invalid while I was working on the take up reel tension. See above entry.
2019-02-07 | Performed QC checks beginning at 13:10 CST.
Conducted site checks. OK.
2019-02-11 | Conducted site checks. OK.
2019-02-15 | Conducted site checks. OK.
2019-02-21 | Performed QC checks beginning at 13:35 CST.
Conducted site checks. OK.
2019-02-25 | Conducted site checks. OK.
2019-03-01 | Conducted site checks. OK.
2019-03-06 | Performed QC checks beginning at 13:25 CST.
Conducted site checks. OK.
2019-03-10 | Hour 02:00 is invalid due to a filter tape pressure alert.
2019-03-12 | Conducted site checks. OK.
2019-03-15 | Trimmed the end of the sample pump line at 09:50 CST.
Conducted site checks. OK.
2019-03-18 | Performed a leak check at 09:40 CST.
As found leak check = 0.5.
Conducted site checks. OK.
2019-03-20 | M. Wolfskill performed a quarterly audit beginning at 08:30 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2019-03-25 | Performed QC checks beginning at 09:10 CST.
Changed out BAM filter tape.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As found leak check = 0.5, As left leak check = 0.5.
2019-03-29 | Conducted site checks. OK.
2019-04-02 | Conducted site checks. OK.
Cleaned PM10 head.
2019-04-05 | Conducted site checks. OK.
2019-04-09 | Conducted site checks. OK.
2019-04-12 | Performed QC checks beginning at 10:25 CST.
Conducted site checks. OK.
2019-04-16 | Conducted site checks. OK.
Inspected enclosure HVAC filter. Found it clean.
2019-04-18 | Conducted site checks. OK.
Hour 18:00 invalid due to an instrument malfunction, delta-pressure error.
2019-04-23 | Conducted site checks. OK.
2019-04-26 | Took the BAM offline at 09:20 CST to check the DeltaCal flow.
It was reading 11.5, instead of the expected 15.0.
Conducted site checks. OK.
2019-04-30 | Performed QC checks beginning at 10:45 CST.
Conducted site checks. OK.
2019-05-03 | Conducted site checks. OK.
2019-05-06 | Conducted site checks. OK.
2019-05-10 | Conducted site checks. OK.
2019-05-16 | Performed QC checks beginning at 09:10 CST.
Conducted site checks. OK.
2019-05-22 | Conducted site checks. OK.
2019-05-24 | Conducted site checks. OK.
2019-05-26 | Recorded negative concentrations of -6 and -15 for hours 21:00 and 22:00 CST, respectively.
2019-05-28 | Performed QC checks beginning at 10:05 CST.
Changed out BAM filter tape.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As found leak check = 0.4, As left leak check = 0.5.
Conducted site checks. OK.
2019-06-04 | Cleaned PM10 sampler head.
Conducted site checks. OK.
2019-06-05 | This site stopped sending data to RAQIS at 08:00 CST. Found the UPS battery backup
had failed. Replaced it with a power strip for now. No data were lost.
2019-06-07 | Power failure during hour 14:00 CST.
2019-06-11 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2019-06-14 | Conducted site checks. OK.
2019-06-18 | Cleaned shelter HVAC filter.
Conducted site checks. OK.
2019-06-21 | Found the main circuit breaker off. The HVAC technician forgot to turn it back on after
completing work on 6/20/2019. As a result, data from 12:00 CST on 6/20/2019 to
09:00 CST on 6/21/2019 are lost. Power was turned off at 12:00 CST to change out
the shelter thermostat. Shelter temperatures have returned to normal.
2019-06-26 | M. Wolfskill performed a quarterly audit beginning at 08:25 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2019-06-28 | Performed QC checks beginning at 08:40 CST.
Conducted site checks. OK.
2019-07-02 | Conducted site checks. OK.
2019-07-05 | Cleaned shelter HVAC filter.
Conducted site checks. OK.
2019-07-09 | Conducted site checks. OK.
2019-07-12 | Performed QC checks beginning at 09:55 CST.
Conducted site checks. OK.
2019-07-16 | Conducted site checks. OK.
2019-07-19 | Conducted site checks. OK.
2019-07-23 | Cleaned PM10 sampler head.
Conducted site checks. OK.
2019-07-26 | Performed QC checks beginning at 09:10 CST.
Conducted site checks. OK.
2019-07-30 | Changed out BAM filter tape at 09:10 CST.
Cleaned BAM nozzle and vane.
Cleaned BAM capstan shaft and pinch roller tires.
As found leak check = 0.5, As left leak check = 0.5.
Conducted site checks. OK.
2019-08-02 | Conducted site checks. OK.
2019-08-06 | Conducted site checks. OK.
2019-08-09 | Performed QC checks beginning at 09:30 CST.
Conducted site checks. OK.
2019-08-13 | Conducted site checks. OK.
2019-08-16 | Conducted site checks. OK.
2019-08-20 | Cleaned shelter HVAC filter.
Conducted site checks. OK.
2019-08-23 | Performed QC checks beginning at 09:10 CST.
Conducted site checks. OK.
2019-08-27 | Conducted site checks. OK.
2019-09-03 | Conducted site checks. OK.
2019-09-05 | Conducted site checks. OK.
2019-09-10 | Performed QC checks beginning at 09:30 CST.
Changed time from 09:33:05 to 09:32 CST.
Conducted site checks. OK.
2019-09-13 | Conducted site checks. OK.
2019-09-19 | M. Wolfskill performed a quarterly audit beginning at 07:30 CST.
Results indicate all systems to be operating within allowable project limits.
Conducted site checks. OK.
2019-09-24 | Conducted site checks. OK.
2019-09-27 | Performed QC checks beginning at 09:15 CST.
Conducted site checks. OK.
2019-09-30 | Conducted site checks. OK.
The BAM display has vertical lines on it and is very hard to read. D. Szatkowski is
sending a replacement. No sign of water in or around the BAM.
2019-10-02 | Hours 02:00 and 03:00 CST are invalid due to a power failure.
2019-10-04 | I was unable to install the new BAM display.
Replaced the existing BAM s/n P21275 with the spare s/n J5560.
Performed QC checks beginning at 12:20 CST.
Conducted site checks. OK.
2019-10-07 | After reviewing RAQIS data over the weekend, it appears the new BAM was not set up
correctly. One of the switches in back was in the wrong position. Checked all of the
BAM settings with D. Szatkowski. We did not lose any data. Data recorded by the BAM
were OK and manually inserted into the database. Changed the filter tape.
2019-10-09 | Conducted site checks. OK.
2019-10-15 | Performed QC checks beginning at 09:20 CST.
Conducted site checks. OK.
2019-10-21 | Conducted site checks. OK.
2019-10-24 | Conducted site checks. OK.
2019-10-28 | Performed QC checks beginning at 12:35 CST.
Conducted site checks. OK.

## ST-Met
2014-10-07 | Conducted site checks. OK.
Found aspirator motor had stopped. Replace with spare. Hour 1400 CST should be
invalid. Checked temperature. Data logger = 19.7, Delta Cal = 19.5 deg C.
D. Szatkowski advised the commucation with this site have been interrupted.
Found a red light on the Cradlepoint. Reset the unit and communications were restored.
The Cradlepoint was reset at 15:00 CST.
2014-10-10 | Conducted site checks. OK.
Found aspirator motor had stopped again. Replaced with spare. The change out
was from 12:30 to 12:45 CST. Since I only changed out the aspirator motor, the ambient
temperature data should be OK. Checked temperature. Data logger = 13.9, Delta cal =
13.8 deg C.
2014-10-13 | Conducted site checks. OK.
Performed WD observation checks @ 10:45. Observed = 90 deg, data logger = 94 deg.
2014-10-16 | Conducted site checks. OK.
2014-10-21 | Conducted site checks. OK.
2014-10-24 | Conducted site checks. OK.
2014-10-28 | Conducted site checks. OK.
2014-10-30 | Conducted site checks. OK.
2014-11-03 | Conducted site checks. OK.
2014-11-05 | Conducted site checks. OK.
2014-11-10 | Conducted site checks. OK.
2014-11-14 | Conducted site checks. OK.
2014-11-17 | Conducted site checks. OK.
2014-11-20 | Conducted site checks. OK.
2014-11-26 | Conducted site checks. OK.
2014-12-01 | Conducted site checks. OK.
2014-12-05 | Conducted site checks. OK.
2014-12-09 | J. Sorkness conducted audits of the meteorological systems from 09:30 to 10:30 CST.
Conducted site checks. OK.
2014-12-12 | Conducted site checks. OK.
2014-12-16 | Conducted site checks. OK.
Changed out data logger from 11:20 to 11:45 CST.
2014-12-18 | Conducted site checks. OK.
2014-12-23 | Conducted site checks. OK.
2014-12-30 | Conducted site checks. OK.
2015-01-02 | Conducted site checks. OK.
2015-01-05 | Conducted site checks. OK.
2015-01-08 | Conducted site checks. OK.
2015-01-12 | Conducted site checks. OK.
2015-01-13 | Conducted site checks. OK.
2015-01-16 | Conducted site checks. OK.
2015-01-20 | Conducted site checks. OK.
2015-01-23 | Conducted site checks. OK.
2015-01-26 | Conducted site checks. OK.
2015-01-29 | Conducted site checks. OK.
Aspirator motor is turning, but sounds weak. Will pickup the spare at the offsite storage
facility and install next week.
2015-02-04 | Conducted site checks. OK.
2015-02-06 | Conducted site checks. OK.
Replaced aspirator from 13:00 to 13:30 CST.
As found temperature on the data logger was -1.7 deg C, as left was -2.1. I will check the
temperature response against the Delta Cal standard during my next visit.
2015-02-10 | Conducted site checks. OK.
2015-02-13 | Conducted site checks. OK.
2015-02-16 | Conducted site checks. OK.
2015-02-19 | Conducted site checks. OK.
2015-02-24 | Conducted site checks. OK.
2015-02-27 | Conducted site checks. OK.
2015-03-02 | Conducted site checks. OK.
2015-03-06 | Conducted site checks. OK.
2015-03-09 | Conducted site checks. OK.
2015-03-11 | Conducted site checks. OK.
2015-03-17 | Conducted site checks. OK.
2015-03-19 | J. Sorkness conducted meteorological audits beginning from 08:15 to 08:40 CST.
2015-03-20 | Conducted site checks. OK.
2015-03-24 | Conducted site checks. OK.
2015-03-27 | Conducted site checks. OK.
2015-03-31 | Conducted site checks. OK.
2015-04-01 | Conducted site checks. OK.
2015-04-08 | Conducted site checks. OK.
2015-04-10 | Conducted site checks. OK.
2015-04-15 | Conducted site checks. OK.
2015-04-17 | Conducted site checks. OK.
2015-04-24 | Conducted site checks. OK.
2015-04-27 | Conducted site checks. OK.
2015-04-30 | Conducted site checks. OK.
2015-05-04 | Conducted site checks. OK.
2015-05-07 | Conducted site checks. OK.
2015-05-11 | Conducted site checks. OK.
2015-05-14 | Conducted site checks. OK.
2015-05-18 | Conducted site checks. OK.
2015-05-21 | Conducted site checks. OK.
2015-05-27 | Conducted site checks. OK.
2015-05-29 | Conducted site checks. OK.
2015-06-01 | Conducted site checks. OK.
2015-06-04 | Conducted site checks. OK.
2015-06-09 | Conducted site checks. OK.
2015-06-12 | Conducted site checks. OK.
2015-06-15 | Conducted site checks.
Temperature reported be reading high for the past several days.
Delta Cal measured temperature to be reading 2.1 deg C high.
This may have been caused by a power issue over the weekend. I found both ST Partisols
to be in "default mode", indicating some kind of power issue.
2015-06-16 | Replaced temperature sensor at 09:45 CST. New probe number is R10882.
2015-06-17 | J. Sorkness performed audit checks beginning at 10:40 to 12:00 CST.
2015-06-18 | Conducted site checks. OK.
2015-06-22 | Conducted site checks. OK.
2015-06-24 | Conducted site checks. OK.
2015-06-30 | Conducted site checks. OK.
2015-07-02 | Conducted site checks. OK.
Tightened the clamps and bolts holding the crossarm to the telephone pole.
Checked alignment of the WD sensor with a survey point (old brick stack).
Compass = 10 deg - 3.6 decl. = 6.4 deg. Datalogger = 10.3 (post tightening).
At 8:30 CST.
2015-07-07 | Conducted site checks. OK.
2015-07-10 | Conducted site checks. OK.
2015-07-14 | Conducted site checks. OK.
2015-07-16 | Conducted site checks. OK.
2015-07-22 | Conducted site checks. OK.
2015-07-24 | Conducted site checks. OK.
2015-07-28 | Conducted site checks. OK.
2015-07-30 | Conducted site checks. OK.
2015-08-05 | Conducted site checks. OK.
2015-08-07 | Conducted site checks. OK.
2015-08-11 | Conducted site checks. OK.
2015-08-12 | J. Sorkness performed an audit of the sensors beginning at 08:00 CST.
The boom bracket bolts and hose clamps were tightened. They did not move.
Conducted site checks. OK.
2015-08-14 | Conducted site checks. OK.
Found door to data logger enclosure open. No water inside. Everything appears to be
working correctly.
2015-08-19 | There was a power failure caused by a lightning strike at 19:00 CST on 8/18.
The data logger continued to record data untill 02:00 this morning, when the battery ran out.
The existing battery only measured 4.4 volts. Replaced with a new one measuring 12.8
volts.
2015-08-21 | Conducted site checks. OK.
2015-08-28 | Conducted site checks. OK.
2015-09-01 | Conducted site checks. OK.
2015-09-04 | Conducted site checks. OK.
2015-09-09 | Conducted site checks. OK.
2015-09-11 | Conducted site checks. OK.
2015-09-15 | Conducted site checks. OK.
2015-09-16 | Conducted site checks. OK.
2015-09-22 | Conducted site checks. OK.
2015-09-25 | Conducted site checks. OK.
2015-09-28 | Conducted site checks. OK.
2015-09-30 | Conducted site checks. OK.
2015-10-02 | Conducted site checks. OK.
2015-10-05 | Conducted site checks. OK.
2015-10-08 | Conducted site checks. OK.
2015-10-13 | Conducted site checks. OK.
2015-10-15 | Conducted site checks. OK.
2015-10-19 | Conducted site checks. OK.
2015-10-22 | Conducted site checks. OK.
2015-10-27 | Conducted site checks. OK.
2015-10-29 | Conducted site checks. OK.
2015-11-02 | Conducted site checks. OK.
Copper grounding strap to the lightning rod has fallen off.
2015-11-04 | Conducted site checks. OK.
2015-11-10 | Conducted site checks. OK.
2015-11-12 | Conducted site checks. OK.
2015-11-16 | Conducted site checks. OK.
Copper grounding strap to the lightning still off.
2015-11-18 | Conducted site checks. OK.
Scheduled audits by J. Sorkness were not performed due to high winds.
2015-11-23 | Conducted site checks. OK.
Copper grounding strap to the lightning still off.
2015-11-25 | Conducted site checks. OK.
2015-11-30 | Conducted site checks. OK.
Copper grounding strap to the lightning rod still off.
2015-12-03 | Conducted audits with Phil Hight from 10:00 to 11:35 CST.
Copper grounding strap reattached to the pole.
Conducted site checks. OK.
2015-12-08 | Conducted site checks. OK.
2015-12-10 | Conducted site checks. OK.
2015-12-14 | Conducted site checks. OK.
2015-12-16 | Conducted site checks. OK.
2015-12-22 | Conducted site checks. OK.
2015-12-29 | Conducted site checks. OK.
2015-12-31 | Conducted site checks. OK.
2016-01-04 | Conducted site checks. OK.
2016-01-06 | Conducted site checks. OK.
2016-01-11 | Conducted site checks. OK.
2016-01-14 | Conducted site checks. OK.
2016-01-19 | Conducted site checks. OK.
2016-01-21 | Conducted site checks. OK.
2016-01-25 | Conducted site checks. OK.
2016-01-27 | Conducted site checks. OK.
2016-02-01 | Conducted site checks. OK.
2016-02-04 | Conducted site checks. OK.
2016-02-08 | Conducted site checks. OK.
2016-02-11 | Conducted site checks. OK.
2016-02-16 | Conducted site checks. OK.
2016-02-19 | Conducted site checks. OK.
2016-02-23 | Conducted site checks. OK.
2016-02-26 | Conducted site checks. OK.
2016-02-29 | Conducted site checks. OK.
2016-03-04 | Conducted site checks. OK.
2016-03-08 | Conducted site checks. OK.
2016-03-11 | Conducted site checks. OK.
2016-03-14 | Conducted site checks. OK.
2016-03-17 | Conducted site checks. OK.
2016-03-23 | J. Sorkness performed system audits.
Wind speed offline at 08:55 CST.
Wind direction offline at 09:15 CST.
2016-03-25 | Conducted site checks. OK.
2016-04-01 | Conducted site checks. OK.
2016-04-04 | Conducted site checks. OK.
2016-04-07 | Conducted site checks. OK.
2016-04-11 | Conducted site checks. OK.
2016-04-14 | Conducted site checks. OK.
2016-04-18 | Conducted site checks. OK.
2016-04-21 | Conducted site checks. OK.
2016-04-25 | Conducted site checks. OK.
2016-04-29 | Conducted site checks. OK.
2016-05-03 | Conducted site checks. OK.
2016-05-05 | Conducted site checks. OK.
2016-05-09 | Conducted site checks. OK.
2016-05-12 | Conducted site checks. OK.
2016-05-16 | Conducted site checks. OK.
2016-05-19 | Conducted site checks. OK.
2016-05-24 | Conducted site checks. OK.
2016-05-27 | Conducted site checks. OK.
2016-05-31 | Conducted site checks. OK.
2016-06-02 | Conducted site checks. OK.
2016-06-07 | J. Sorkness performed system audits.
Wind speed, wind direction, AT, and BP were offline at 07:55 CST.
Wind speed, temperature, and BP were found to be operating within allowable audit limits.
During the WD audit, the existing sensor response was +3 degrees high from the 270 to
330 degree sectors. These results did not meet audit tolerences. After consulting with
J. Clarke, it was determined to change out the WD sensor. The original WD sensor was
replaced with sensor S/N M3325. The S/N of the old WD sensor was bleached out by the
sun and not readable. The declination of the old WD sensor was 3.5 degrees, and the new
sensors was 4.2 degrees, both within allowable project audit limits.
2016-06-10 | Conducted site checks. OK.
2016-06-14 | Conducted site checks. OK.
2016-06-20 | Conducted site checks. OK.
2016-06-23 | Conducted site checks. OK.
2016-06-27 | Conducted site checks. OK.
2016-06-30 | Conducted site checks. OK.
2016-07-02 | Apparent calm conditions during hours 01:00-02:00 CST.
2016-07-06 | Conducted site checks. OK.
2016-07-08 | Conducted site checks. OK.
2016-07-11 | Conducted site checks. OK.
2016-07-14 | Conducted site checks. OK.
2016-07-19 | Conducted site checks. OK.
2016-07-22 | Conducted site checks. OK.
2016-07-26 | Conducted site checks. OK.
Apparent calm conditions during hours 02:00-0:400 CST.
2016-07-28 | Conducted site checks. OK.
Apparent calm conditions during hours 03:00-04:00 CST.
2016-08-02 | Conducted site checks. OK.
2016-08-05 | Conducted site checks. OK.
2016-08-09 | Conducted site checks. OK.
2016-08-12 | Conducted site checks. OK.
2016-08-15 | Conducted site checks. OK.
2016-08-18 | Conducted site checks. OK.
2016-08-22 | Conducted site checks. OK.
2016-08-26 | Conducted site checks. OK.
2016-09-01 | J. Sorkness performed system audits.
All meteorological parameters were found to be operating within allowable project limits.
The WD sensor was properly aligned.
The audits began at 08:00 CST and ended at 08:45 CST.
2016-09-06 | Conducted site checks. OK.
2016-09-09 | Conducted site checks. OK.
2016-09-13 | Conducted site checks. OK.
2016-09-16 | Conducted site checks. OK.
2016-09-19 | Conducted site checks. OK.
2016-09-22 | Conducted site checks. OK.
2016-10-01 | Conducted site checks. OK.
2016-10-04 | Conducted site checks. OK.
2016-10-07 | Conducted site checks. OK.
2016-10-11 | Conducted site checks. OK.
2016-10-14 | Conducted site checks. OK.
2016-10-17 | Conducted site checks. OK.
2016-10-20 | Conducted site checks. OK.
2016-10-25 | Conducted site checks. OK.
2016-10-28 | Conducted site checks. OK.
2016-11-02 | Conducted site checks. OK.
2016-11-04 | Conducted site checks. OK.
2016-11-08 | Conducted site checks. OK.
2016-11-11 | Conducted site checks. OK.
2016-11-15 | Conducted site checks. OK.
2016-11-18 | Conducted site checks. OK.
2016-11-22 | Conducted site checks. OK.
2016-11-25 | Conducted site checks. OK.
2016-11-30 | Conducted site checks. OK.
2016-12-02 | J. Sorkness performed system audits.
All meteorological parameters were found to be operating within allowable project limits.
The WD sensor was properly aligned.
The audits began at 08:50 CST and ended at 09:30 CST.
2016-12-06 | Conducted site checks. OK.
2016-12-09 | Conducted site checks. OK.
2016-12-13 | Conducted site checks. OK.
2016-12-16 | Conducted site checks. OK.
2016-12-20 | Conducted site checks. OK.
2016-12-22 | Conducted site checks. OK.
2016-12-28 | Conducted site checks. OK.
2016-12-30 | Conducted site checks. OK.
2017-01-03 | Conducted site checks. OK.
2017-01-06 | Conducted site checks. OK.
2017-01-09 | Conducted site checks. OK.
2017-01-13 | Conducted site checks. OK.
2017-01-17 | Conducted site checks. OK.
2017-01-23 | Conducted site checks. OK.
2017-01-27 | Conducted site checks. OK.
2017-01-31 | Conducted site checks. OK.
2017-02-06 | Conducted site checks. OK.
2017-02-13 | Conducted site checks. OK.
2017-02-17 | Conducted site checks. OK.
2017-02-23 | Conducted site checks. OK.
2017-02-27 | Conducted site checks. OK.
2017-03-03 | Conducted site checks. OK.
2017-03-08 | Conducted site checks. OK.
2017-03-16 | Conducted site checks. OK.
2017-03-20 | Conducted site checks. OK.
2017-03-24 | J. Sorkness performed system audits.
All meteorological parameters were found to be operating within allowable project limits.
The WD sensor was properly aligned.
The audits began at 08:00 CST and ended at 08:50 CST.
2017-03-27 | Conducted site checks. OK.
2017-03-31 | Conducted site checks. OK.
2017-04-07 | Conducted site checks. OK.
2017-04-10 | Conducted site checks. OK.
2017-04-13 | Conducted site checks. OK.
2017-04-18 | Conducted site checks. OK.
2017-04-21 | Conducted site checks. OK.
2017-04-25 | Conducted site checks. OK.
2017-04-28 | Conducted site checks. OK.
2017-05-05 | Conducted site checks. OK.
2017-05-12 | Conducted site checks. OK.
2017-05-16 | Conducted site checks. OK.
2017-05-19 | Conducted site checks. OK.
2017-05-23 | Conducted site checks. OK.
2017-05-26 | Conducted site checks. OK.
2017-05-30 | Conducted site checks. OK.
2017-06-05 | Conducted site checks. OK.
2017-06-08 | Conducted site checks. OK.
2017-06-13 | Conducted site checks. OK.
2017-06-14 | A storm blew out the fuse to the meteorological station. The battery backup lasted until
17:00 CST.
2017-06-15 | The fuse was replaced at 11:20 CST by the local power company. Valid data collection
resumed. The backup battery was replaced. It was reading only 2.7 volts DC.
2017-06-20 | Conducted site checks. OK.
2017-06-23 | Conducted site checks. OK.
2017-06-29 | J. Sorkness performed system audits.
All meteorological parameters were found to be operating within allowable project limits.
The WD sensor was properly aligned.
The audits began at 08:00 CST and ended at 08:55 CST.
2017-07-05 | Conducted site checks. OK.
2017-07-11 | Conducted site checks. OK.
2017-07-14 | Conducted site checks. OK.
2017-07-18 | Conducted site checks. OK.
2017-07-21 | Conducted site checks. OK.
2017-07-24 | Conducted site checks. OK.
2017-07-27 | Conducted site checks. OK.
2017-07-31 | Conducted site checks. OK.
2017-08-04 | Conducted site checks. OK.
2017-08-08 | Conducted site checks. OK.
2017-08-11 | Conducted site checks. OK.
2017-08-15 | Conducted site checks. OK.
2017-08-16 | Conducted site checks. OK.
2017-08-22 | Conducted site checks. OK.
2017-08-25 | Conducted site checks. OK.
2017-08-29 | Conducted site checks. OK.
2017-08-31 | Conducted site checks. OK.
2017-09-05 | Conducted site checks. OK.
2017-09-08 | J. Sorkness performed system audits.
All meteorological parameters were found to be operating within allowable project limits.
The WD sensor was properly aligned.
The audits began at 07:50 CST and ended at 08:40 CST.
2017-09-12 | Conducted site checks. OK.
2017-09-15 | Conducted site checks. OK.
2017-09-19 | Conducted site checks. OK.
2017-09-22 | Conducted site checks. OK.
2017-10-02 | Conducted site checks. OK.
2017-10-05 | Conducted site checks. OK.
2017-10-09 | Conducted site checks. OK.
2017-10-12 | Conducted site checks. OK.
2017-10-16 | Conducted site checks. OK.
2017-10-20 | Conducted site checks. OK.
2017-10-24 | Conducted site checks. OK.
2017-10-27 | Conducted site checks. OK.
2017-11-06 | Conducted site checks. OK.
2017-11-10 | Conducted site checks. OK.
2017-11-14 | Conducted site checks. OK.
2017-11-17 | Conducted site checks. OK.
2017-11-24 | Conducted site checks. OK.
2017-11-27 | Conducted site checks. OK.
2017-11-30 | J. Sorkness performed system audits.
All meteorological parameters were found to be operating within allowable project limits.
The WD sensor was properly aligned.
The audits began at 08:50 CST and ended at 09:20 CST.
2017-12-05 | Conducted site checks. OK.
2017-12-12 | Conducted site checks. OK.
2017-12-15 | Conducted site checks. OK.
2017-12-19 | Conducted site checks. OK.
2017-12-22 | Conducted site checks. OK.
2017-12-26 | Meteorological system backup battery failed due to bitter cold temperatures.
Replaced battery and charger. Data from 12/25/17 at 23:00 to 12/26/17 at 11:00 lost.
2017-12-29 | Conducted site checks. OK.
2018-01-03 | Conducted site checks. OK.
2018-01-09 | Conducted site checks. OK.
2018-01-12 | Conducted site checks. OK.
2018-01-17 | Conducted site checks. OK.
2018-01-19 | Conducted site checks. OK.
2018-01-22 | Conducted site checks. OK.
2018-01-25 | Conducted site checks. OK.
2018-01-29 | Conducted site checks. OK.
2018-01-30 | Hour 06:00 appears to be calm conditions. Checked with NWS stations in the area and
they were recording calm conditions as well.
2018-02-02 | Conducted site checks. OK.
2018-02-06 | Conducted site checks. OK.
Hours 2/6/18 19:00 to 2/7/18 00:00 CST appear to be calm conditions.
2018-02-13 | Conducted site checks. OK.
2018-02-15 | Hour 19:00 CST appears to be calm conditions.
2018-02-16 | Conducted site checks. OK.
2018-02-20 | Conducted site checks. OK.
2018-02-23 | Conducted site checks. OK.
Hour 20:00 CST appears to be calm conditions.
2018-02-26 | Conducted site checks. OK.
Hours 03:00 to 05:00 CST appear to be calm conditions.
2018-03-01 | Conducted site checks. OK.
2018-03-02 | Hours 19:00 to 3/3/18 at 01:00 CST appear to be calm conditions.
2018-03-06 | Conducted site checks. OK.
2018-03-09 | Conducted site checks. OK.
2018-03-10 | Hours 05:00 to 07:00 CST appear to be calm conditions.
2018-03-11 | Hour 23:00 CST appears to be calm conditions.
2018-03-13 | Conducted site checks. OK.
2018-03-16 | Conducted site checks. OK.
2018-03-17 | Hours 21:00 to 3/18/18 at 05:00 CST appear to be calm conditions.
2018-03-18 | Hours 19:00 to 21:00 CST appear to be calm conditions.
2018-03-20 | Conducted site checks. OK.
2018-03-23 | Conducted site checks. OK.
2018-03-27 | Conducted site checks. OK.
2018-03-29 | J. Sorkness performed system audits.
All meteorological parameters were found to be operating within allowable project limits.
The WD sensor was properly aligned.
The audits began at 07:45 CST and ended at 08:30 CST.
2018-04-03 | Conducted site checks. OK.
2018-04-06 | Conducted site checks. OK.
2018-04-07 | Hours 21:00 to 4/8/18 at 01:00 CST appear to be calm conditions.
2018-04-10 | Conducted site checks. OK.
2018-04-13 | Conducted site checks. OK.
2018-04-16 | WS and WD sensors were frozen beginning at 23:00 on 4/15/18. Both sensors were
moving when I arrived on site at 11:30. With the help of a JLG and KCBX operator,
I inspected the sensors and reinstalled the ice shields.
2018-04-19 | Conducted site checks. OK.
2018-04-20 | Hour 02:00 appears to be calm conditions.
2018-04-21 | Hour 01:00 appears to be calm conditions.
2018-04-23 | Conducted site checks. OK.
2018-04-25 | Hours 21:00 to 4/26/18 at 03:00 CST appear to be calm conditions.
2018-04-26 | Conducted site checks. OK.
2018-05-01 | Conducted site checks. OK.
2018-05-04 | Conducted site checks. OK.
2018-05-05 | Hours 00:00 to 03:00 CST appear to be calm conditions.
2018-05-06 | Hours 00:00 to 03:00 CST appear to be calm conditions.
2018-05-11 | Conducted site checks. OK.
2018-05-13 | Hour 22:00 CST appears to be calm conditions.
2018-05-14 | Hour 20:00 CST appears to be calm conditions.
2018-05-16 | Hours 01:00 to 03:00 CST appear to be calm conditions.
Hour 19:00 to hour 04:00 CST on 5/17 appear to be calm conditions.
2018-05-18 | Conducted site checks. OK.
D. Szatkowski observed more calm conditions than ususal. Thinks the WS bearings may
be starting to wear out. Replaced WS s/n P15943 with WS s/n W3522 at 09:50 CST.
2018-05-22 | Conducted site checks. OK.
2018-05-25 | Conducted site checks. OK.
2018-05-29 | Conducted site checks. OK.
2018-06-01 | Conducted site checks. OK.
2018-06-05 | Conducted site checks. OK.
2018-06-08 | Conducted site checks. OK.
2018-06-11 | Conducted site checks. OK.
Replaced WD sensor s/n M3325 with sensor s/n R19037.
2018-06-14 | Conducted site checks. OK.
2018-06-18 | Conducted site checks. OK.
2018-06-21 | Conducted site checks. OK.
2018-06-25 | Conducted site checks. OK.
2018-06-28 | J. Sorkness performed system audits.
All meteorological parameters were found to be operating within allowable project limits
except for WD. The WD sensor was reading seven degrees low, and was adjusted.
The audits began at 08:00 CST and ended at 09:00 CST.
2018-07-02 | Conducted site checks. OK.
2018-07-05 | Conducted site checks. OK.
2018-07-10 | Conducted site checks. OK.
2018-07-13 | Conducted site checks. OK.
2018-07-16 | Conducted site checks. OK.
2018-07-19 | Conducted site checks. OK.
2018-07-24 | Conducted site checks. OK.
2018-07-27 | Conducted site checks. OK.
2018-07-30 | Conducted site checks. OK.
2018-08-03 | Conducted site checks. OK.
2018-08-07 | Conducted site checks. OK.
2018-08-10 | Conducted site checks. OK.
2018-08-14 | Conducted site checks. OK.
2018-08-17 | Conducted site checks. OK.
2018-08-23 | Conducted site checks. OK.
2018-08-28 | Met system had not been polled by RAQIS since 8/27/18 at 11:00 CST. Found all power
was on. Power cycled all systems. D. Szatkowski reported that the met system is back
online and no data were lost.
2018-08-30 | Conducted site checks. OK.
2018-09-04 | Conducted site checks. OK.
2018-09-07 | Conducted site checks. OK.
2018-09-11 | Conducted site checks. OK.
2018-09-14 | Conducted site checks. OK.
2018-09-18 | Conducted site checks. OK.
2018-09-21 | Conducted site checks. OK.
2018-09-24 | Met system had not been polled by RAQIS since 9/23/18 at 21:00 CST. Found all power
was on. Power cycled all systems. D. Szatkowski reported that the met system is back
online and no data were lost.
2018-09-27 | J. Sorkness performed system audits.
All meteorological parameters were found to be operating within allowable project limits
except for WD. The WD sensor failed the linearity test. Replaced the existing sensor
with the spare. Results were worse. D. Szatkowski will send a replacement WD sensor
The existing WD sensor was aligned correctly.
The audits began at 08:00 CST and ended at 09:30 CST.
2018-10-02 | Conducted site checks. OK.
2018-10-05 | Conducted site checks. OK.
2018-10-12 | Conducted site checks. OK.
2018-10-16 | WD sensor s/n R19037 was replaced with WD sensor s/n X16490 at 10:00 CST.
Checked N, E, S, and W sensor responses with D. Szatkowski.
Conducted site checks. OK.
2018-10-18 | Met data stopped being recorded by RAQIS at 18:00 CST.
2018-10-19 | Found a loose power wire to the data logger and repaired it at 08:45 CST.
Because of this, met data were lost from 10/18/2018 18:00 to 10/19/2018 09:00 CST.
2018-10-23 | Conducted site checks. OK.
2018-10-26 | Conducted site checks. OK.
2018-10-30 | Conducted site checks. OK.
2018-11-02 | Conducted site checks. OK.
2018-11-06 | RAQIS stopped recording met data. Rebooted the data logger and data collection resumed.
Conducted site checks. OK.
2018-11-12 | Conducted site checks. OK.
2018-11-16 | Conducted site checks. OK.
2018-11-27 | Conducted site checks. OK.
2018-11-30 | Conducted site checks. OK.
2018-12-03 | RAQIS stopped recording met data. Rebooted the data logger and data collection resumed.
Conducted site checks. OK.
2018-12-06 | J. Sorkness performed system audits.
All meteorological parameters were found to be operating within allowable project limits.
The WD sensor was aligned correctly.
The audits began at 08:30 CST and ended at 09:40 CST.
2018-12-11 | Conducted site checks. OK.
2018-12-14 | Conducted site checks. OK.
2018-12-18 | Conducted site checks. OK.
2018-12-21 | Conducted site checks. OK.
2018-12-28 | Conducted site checks. OK.
2018-12-31 | RAQIS stopped recording met data. Rebooted the data logger and data collection resumed.
Conducted site checks. OK.
2019-01-08 | Conducted site checks. OK.
2019-01-11 | Conducted site checks. OK.
2019-01-15 | Conducted site checks. OK.
2019-01-18 | Conducted site checks. OK.
2019-01-24 | Conducted site checks. OK.
2019-01-25 | RAQIS stopped recording data at 06:00 CST. Visited the site and found the power light
on the charger off. Replaced the charger and battery, and data recording resumed
at 11:00 CST. Hours 06:00 to 10:00 are invalid due to the power failure.
2019-02-01 | Conducted site checks. OK.
2019-02-07 | Conducted site checks. OK.
2019-02-11 | Conducted site checks. OK.
2019-02-15 | Conducted site checks. OK.
2019-02-18 | Trying to touble shoot communications issues by replacing the NL201.
2019-02-21 | Conducted site checks. OK.
2019-02-25 | Conducted site checks. OK.
RAQIS stopped recording met data. Rebooted the data logger and data collection resumed.
Conducted site checks. OK.
2019-03-01 | Conducted site checks. OK.
2019-03-06 | Conducted site checks. OK.
Replaced the data logger. Hour 11:00 CST is invalid for all sensors.
2019-03-12 | Conducted site checks. OK.
2019-03-15 | Conducted site checks. OK.
2019-03-18 | Conducted site checks. OK.
2019-03-20 | M. Wolfskill performed system audits.
All meteorological parameters were found to be operating within allowable project limits.
The WD sensor was aligned correctly.
The audits began at 14:00 CST and ended at 14:50 CST.
2019-03-25 | Conducted site checks. OK.
2019-03-29 | Conducted site checks. OK.
2019-04-02 | Conducted site checks. OK.
2019-04-05 | Conducted site checks. OK.
2019-04-09 | Conducted site checks. OK.
2019-04-12 | Conducted site checks. OK.
2019-04-16 | Conducted site checks. OK.
2019-04-18 | Conducted site checks. OK.
2019-04-23 | Conducted site checks. OK.
2019-04-26 | Conducted site checks. OK.
2019-04-30 | Conducted site checks. OK.
2019-05-03 | Conducted site checks. OK.
2019-05-06 | Conducted site checks. OK.
2019-05-10 | Conducted site checks. OK.
2019-05-16 | Conducted site checks. OK.
2019-05-22 | Conducted site checks. OK.
2019-05-24 | Conducted site checks. OK.
2019-05-28 | Conducted site checks. OK.
2019-06-04 | Conducted site checks. OK.
2019-06-05 | RAQIS stopped recording data at 13:00 CST on 6/4/2019.
Visited the site and found the power light on the charger off.
Replaced the charger and battery, and data recording resumed at 11:00 CST.
Data from 13:00 CST on 6/4/2019 to 10:00 CST on 6/5/2019 are invalid due to
the power failure.
2019-06-11 | Conducted site checks. OK.
2019-06-14 | Conducted site checks. OK.
2019-06-18 | Conducted site checks. OK.
2019-06-21 | Conducted site checks. OK.
2019-06-26 | M. Wolfskill performed system audits.
All meteorological parameters were found to be operating within allowable project limits.
The WD sensor was aligned correctly.
The audits began at 08:00 CST and ended at 08:40 CST.
2019-06-28 | Conducted site checks. OK.
2019-07-02 | Conducted site checks. OK.
2019-07-05 | Conducted site checks. OK.
2019-07-09 | Conducted site checks. OK.
2019-07-12 | Conducted site checks. OK.
2019-07-16 | Conducted site checks. OK.
2019-07-19 | Conducted site checks. OK.
2019-07-23 | Conducted site checks. OK.
2019-07-26 | Conducted site checks. OK.
2019-07-30 | Conducted site checks. OK.
2019-08-02 | Conducted site checks. OK.
2019-08-06 | Conducted site checks. OK.
2019-08-09 | Conducted site checks. OK.
2019-08-13 | Conducted site checks. OK.
2019-08-16 | Conducted site checks. OK.
2019-08-20 | Conducted site checks. OK.
2019-08-23 | Conducted site checks. OK.
2019-08-27 | Conducted site checks. OK.
2019-09-03 | Conducted site checks. OK.
2019-09-05 | Conducted site checks. OK.
2019-09-10 | Conducted site checks. OK.
2019-09-13 | Conducted site checks. OK.
2019-09-19 | M. Wolfskill performed system audits.
All meteorological parameters were found to be operating within allowable project limits.
The WD sensor was aligned correctly.
The audits began at 08:30 CST and ended at 09:00 CST.
2019-09-20 | Inspected the meteorological system after reviewing the wind data and finding hours 03:00
and 04:00 recording all zeros or missing. Everything seems to be operating normally.
Checked with B. Stringer, and he indicated the wind instruments look good from his end.
The area experienced very heavy rains during that period.
2019-09-24 | Conducted site checks. OK.
2019-09-27 | Conducted site checks. OK.
2019-09-30 | Conducted site checks. OK.
2019-10-04 | Conducted site checks. OK.
2019-10-07 | Conducted site checks. OK.
2019-10-09 | Conducted site checks. OK.
2019-10-15 | Conducted site checks. OK.
2019-10-21 | Conducted site checks. OK.
2019-10-24 | Conducted site checks. OK.
2019-10-28 | Conducted site checks. OK.