## MetBAM Checklist South
KCBX TERMINALS CO SOUTH |  | SITE VISIT DATE
| 2020-05-01 | 2020-05-06 | 2020-05-12 | 2020-05-15 | 2020-05-19 | 2020-05-22 | 2020-05-28
Meteorological Station
Arrival time: | 8:50 | 9:00 | 9:40 | 9:30 | 10:35 | 8:50 | 9:35
Is the temperature aspirator fan running? | Y | Y | Y | Y | Y | Y | Y
Are all signal cables secured and intact? | Y | Y | Y | Y | Y | Y | Y
Are wind vane/ anemometer cups intact and operational? | Y | Y | Y | Y | Y | Y | Y
Is datalogger date/time correct within 60 sec? | NA | NA | NA | NA | NA | NA | NA
Are conditions shown on logger consistent with ambient observations? | NA | NA | NA | NA | NA | NA | NA
BAM-1020 PM10 Monitors
Arrival time: | 8:50 | 9:00 | 9:40 | 9:30 | 10:35 | 8:50 | 9:35
Any status error flags showing? | N | N | N | N | N | N | N
Are all monitor dates/times correct within 60 sec? | Y | N | Y | Y | Y | N | Y
Are all filter transport systems operational? | Y | Y | Y | Y | Y | Y | Y
Are all tape reel cover knobs secure? | Y | Y | Y | Y | Y | Y | Y
Are hourly data being collected? | Y | Y | Y | Y | Y | Y | Y
Is sampler inlet and downtube clean? | Y | Y | Y | Y | Y | Y | Y
Is water collection bottle clean? | Y | Y | Y | Y | Y | Y | Y
Are O-rings in good condition? | Y | Y | Y | Y | Y | Y | Y
Comments section
5/6, Replaced the ST-NW BAM.
5/6, Performed QC checks on the ST-NW BAM.
5/12, Perfomed BAM QC checks.
5/15, Replaced a leaky quick connect on the ST-NW BAM pump.
5/15, Performed QC checks on the ST-NW BAM.
5/19, Changed time on the ST-SE BAM from 10:52:05 to 10:51 CST.
5/22, Performed QC checks.
5/24, Replaced the ST-NW BAM.
5/24, Perfromed QC checks on the ST-NW BAM.
5/29, Replaced the UPS powering the coms, with a power strip
on the ST-SW BAM enclosure.
Site ID List: |  | MAINTENANCE LIST FOR METEOROLOGICAL STATION
ST-NW = Northwest BAM site |  | Each Visit: | Inspect sensors and cables for signs of deterioration or damage and replace as needed
ST-N = North BAM site |  | Monthly: | Inspect temperature aspirator and remove dust accumulation
ST-NE = Northeast BAM site (Partisol samplers) |  |  | Inspect tower for plumb, and verify that all fasteners are intact and tight
ST-M = Meteorological Monitoring Site |  | Biannually: | Perform calibration check on sensors
ST-CE = Center East BAM site |  |  | Check wind sensor bearings
ST-SW = Southwest BAM site |  | Annually: | Replace wind sensor bearings
MAINTENANCE LIST FOR BAM-1020
|  | Each Visit: | Inspect filter tape and replace when needed (~2 months/roll; run SELF-TEST function in the TAPE menu).
Inspect filter tape for pinholes and irregular PM deposits
|  | Monthly: | Perform flow rate and leak check
Clean BAM nozzle and vane
Clean inlet hardware, capstan shaft and pinch roller
|  | Quarterly: | Verify BAM settings file
Flow system calibration
Disassemble and clean inlet
|  | Biannually: | Check pump muffler and tubing and replace as needed
Test filter RH and Temperature sensors
Test smart heater function
|  | Annually: | Perform zero background test
Clean internal debris filter
Remove and check membrane span foil
Beta detector count rate and dark count test
Rebuild vacuum pump
Replace nozzle o-ring