## MetBAM Checklist South
KCBX TERMINALS CO SOUTH |  | SITE VISIT DATE
| 2020-11-03 | 2020-11-06 | 2020-11-10 | 2020-11-13 | 2020-11-17 | 2020-11-20 | 2020-11-25 | 2020-11-30
Meteorological Station
Arrival time: | 10:15 | 10:10 | 10:35 | 10:05 | 10:15 | 10:30 | 10:15 | 10:10
Is the temperature aspirator fan running? | bar: 2020-11-03 to 2020-11-30
Are all signal cables secured and intact? | bar: 2020-11-03 to 2020-11-30
Are wind vane/ anemometer cups intact and operational? | bar: 2020-11-03 to 2020-11-30
Is datalogger date/time correct within 60 sec? | bar: 2020-11-03 to 2020-11-30
Are conditions shown on logger consistent with ambient observations? | bar: 2020-11-03 to 2020-11-30
BAM-1020 PM10 Monitors
Arrival time: | 10:15 | 10:10 | 10:35 | 10:05 | 10:15 | 10:30 | 10:15 | 10:10
Any status error flags showing? | bar: 2020-11-03 to 2020-11-30
Are all monitor dates/times correct within 60 sec? | bar: 2020-11-03 to 2020-11-30
Are all filter transport systems operational? | bar: 2020-11-03 to 2020-11-30
Are all tape reel cover knobs secure? | bar: 2020-11-03 to 2020-11-30
Are hourly data being collected? | bar: 2020-11-03 to 2020-11-30
Is sampler inlet and downtube clean? | bar: 2020-11-03 to 2020-11-30
Is water collection bottle clean? | bar: 2020-11-03 to 2020-11-30
Are O-rings in good condition? | bar: 2020-11-03 to 2020-11-30
Comments section
11/3, Changed out the ST-N BAM filter tape.
11/3, Changed time on the ST-SE BAM from 10:45:05 to 10:44 CST.
11/10, Cleaned shelter HVAC filters.
11/13, Changed out the ST-NW filter tape.
11/17, Performed BAM QC checks.
11/20, Changed out the ST-NE and ST-SW BAM filter tapes.
11/25, Checked out a flow issue with the ST-SW BAM.
11/30, Performed BAM QC checks.
11/30, Changed time on the ST-SW BAM from 10:53:05 to 10:52 CST.
Site ID List: | 2020-11-06: MAINTENANCE LIST FOR METEOROLOGICAL STATION
ST-NW = Northwest BAM site | 2020-11-06: Each Visit:; 2020-11-10: Inspect sensors and cables for signs of deterioration or damage and replace as needed
ST-N = North BAM site | 2020-11-06: Monthly:; 2020-11-10: Inspect temperature aspirator and remove dust accumulation
ST-NE = Northeast BAM site (Partisol samplers) | 2020-11-10: Inspect tower for plumb, and verify that all fasteners are intact and tight
ST-M = Meteorological Monitoring Site | 2020-11-06: Biannually:; 2020-11-10: Perform calibration check on sensors
ST-CE = Center East BAM site | 2020-11-10: Check wind sensor bearings
ST-SW = Southwest BAM site | 2020-11-06: Annually:; 2020-11-10: Replace wind sensor bearings
MAINTENANCE LIST FOR BAM-1020
|  | Each Visit: | Inspect filter tape and replace when needed (~2 months/roll; run SELF-TEST function in the TAPE menu).
Inspect filter tape for pinholes and irregular PM deposits
|  | Monthly: | Perform flow rate and leak check
Clean BAM nozzle and vane
Clean inlet hardware, capstan shaft and pinch roller
|  | Quarterly: | Verify BAM settings file
Flow system calibration
Disassemble and clean inlet
|  | Biannually: | Check pump muffler and tubing and replace as needed
Test filter RH and Temperature sensors
Test smart heater function
|  | Annually: | Perform zero background test
Clean internal debris filter
Remove and check membrane span foil
Beta detector count rate and dark count test
Rebuild vacuum pump
Replace nozzle o-ring