## MetBAM Checklist South
KCBX TERMINALS CO SOUTH |  | SITE VISIT DATE
| 2021-08-03 | 2021-08-06 | 2021-08-10 | 2021-08-13 | 2021-08-17 | 2021-08-20 | 2021-08-24 | 2021-08-27 | 2021-08-31
Meteorological Station
Arrival time: | 8:50 | 7:30 | 9:00 | 9:05 | 9:00 | 9:45 | 9:00 | 8:40 | 7:15
Is the temperature aspirator fan running? | bar: 2021-08-03 to 2021-08-31
Are all signal cables secured and intact? | bar: 2021-08-03 to 2021-08-31
Are wind vane/ anemometer cups intact and operational? | bar: 2021-08-03 to 2021-08-31
Is datalogger date/time correct within 60 sec? | bar: 2021-08-03 to 2021-08-31
Are conditions shown on logger consistent with ambient observations? | bar: 2021-08-03 to 2021-08-31
BAM-1020 PM10 Monitors
Arrival time: | 8:50 | 7:30 | 9:00 | 9:05 | 9:00 | 9:45 | 9:00 | 8:40 | 7:15
Any status error flags showing? | bar: 2021-08-03 to 2021-08-31
Are all monitor dates/times correct within 60 sec? | bar: 2021-08-03 to 2021-08-31
Are all filter transport systems operational? | bar: 2021-08-03 to 2021-08-31
Are all tape reel cover knobs secure? | bar: 2021-08-03 to 2021-08-31
Are hourly data being collected? | bar: 2021-08-03 to 2021-08-31
Is sampler inlet and downtube clean? | bar: 2021-08-03 to 2021-08-31
Is water collection bottle clean? | bar: 2021-08-03 to 2021-08-31
Are O-rings in good condition? | bar: 2021-08-03 to 2021-08-31
8/3, Changed out ST-N BAM filter tape.
8/6, Replaced ST-SE sample pump.
8/6, Replaced ST-NE HVAC system.
8/10, Performed BAM QC checks.
8/13, Changed time on the ST-NW BAM from 09:22:50 to 09:24 CST.
8/20, Changed out the ST-NE, ST-SE, and ST-SW BAM filter tapes.
8/24, Performed BAM QC checks.
8/27, Changed time on the ST-NW BAM from 08:45:50 to 08:47 CST.
8/31, M. Hale performed audits.
Site ID List: | 2021-08-06: MAINTENANCE LIST FOR METEOROLOGICAL STATION
ST-NW = Northwest BAM site | 2021-08-06: Each Visit:; 2021-08-10: Inspect sensors and cables for signs of deterioration or damage and replace as needed
ST-N = North BAM site | 2021-08-06: Monthly:; 2021-08-10: Inspect temperature aspirator and remove dust accumulation
ST-NE = Northeast BAM site (Partisol samplers) | 2021-08-10: Inspect tower for plumb, and verify that all fasteners are intact and tight
ST-M = Meteorological Monitoring Site | 2021-08-06: Biannually:; 2021-08-10: Perform calibration check on sensors
ST-CE = Center East BAM site | 2021-08-10: Check wind sensor bearings
ST-SW = Southwest BAM site | 2021-08-06: Annually:; 2021-08-10: Replace wind sensor bearings
MAINTENANCE LIST FOR BAM-1020
|  | Each Visit: | Inspect filter tape and replace when needed (~2 months/roll; run SELF-TEST function in the TAPE menu).
Inspect filter tape for pinholes and irregular PM deposits
|  | Monthly: | Perform flow rate and leak check
Clean BAM nozzle and vane
Clean inlet hardware, capstan shaft and pinch roller
|  | Quarterly: | Verify BAM settings file
Flow system calibration
Disassemble and clean inlet
|  | Biannually: | Check pump muffler and tubing and replace as needed
Test filter RH and Temperature sensors
Test smart heater function
|  | Annually: | Perform zero background test
Clean internal debris filter
Remove and check membrane span foil
Beta detector count rate and dark count test
Rebuild vacuum pump
Replace nozzle o-ring