## Recommended Maintenance List
Each Visit: | Inspect filter tape and replace when needed (~2 months/roll; run SELF-TEST function in the TAPE menu).
Inspect filter tape for pinholes and irregular PM deposits
Monthly: | Perform flow rate and leak check
Clean BAM nozzle and vane
Clean inlet hardware, capstan shaft and pinch roller
Quarterly: | Verify BAM settings file
Flow system Calibration
Disassemble and clean inlet
Biannually: | Check pump muffler and tubing and replace as needed
Test filter RH and Temperature sensors
Test smart heater function
Annually: | Perform zero background test
Clean internal debris filter
Remove and check membrane span foil
Beta detector count rate and dark count test
Rebuild vacuum pump
Replace nozzle o-ring

## Equipment Info
KCBX ST-SW - Equipment Information
Current Delta Cal Info
Serial Number:
Expiration Date:
put in service
Current BAM Serial Numbers
| Site | PM-2.5
Site 1
| The information above is current. The comments below provide historical information on what equipment was previously used. |  | being used previously
Not applicable as of 12/12/14.

## ST-NE
KCBX ST-NE Operator Checklist |  | 10 weeks
Arrival Date: (to insert new dates selet Column C. - Right click - insert 1 left) | 2022-08-25 | 2022-09-02 | 2022-09-08 | 2022-09-16 | 2022-09-21 | 2022-09-27 | 2022-10-07 | 2022-10-14 | 2022-10-19 | 2022-11-02 | 2022-11-14 | 2022-11-28 | 2022-12-13 | 2022-12-28 | 2023-01-05 | 2023-01-13 | 2023-01-20 | 2023-01-27 | 2023-02-06 | 2023-02-24 | 2023-03-10 | 2023-03-30 | 2023-04-14 | 2023-04-21 | 2015-10-28 | 2015-04-06 | 2015-04-14 | 2015-04-21 | 2015-05-06 | 2015-05-20 | 2015-06-03 | 2015-06-19 | 2015-07-06 | 2015-07-21 | 2015-08-04 | 2015-08-17 | 2015-09-03 | 2015-09-16 | 2015-10-01
Arrival Time: | 11:30 | 8:10 | 9:45 | 9:15 | 9:25 | 10:25 | 9:00 | 9:30 | 11:30 | 12:00 | 11:00 | 9:05 | 11:00 | 11:20 | 12:00 | 11:15 | 10:00 | 9:00 | 11:00 | 10:00 | 11:00 | 11:40 | 11:00 | 10:40 | 1111 |  |  | 1125 | 9:05 | 11:14 | 14:00 | 1245 | 1303 | 1427 | 1320 | 1255 | 1136 | 1058 | 13:40
BAM-1020 PM10 Monitor
Is it time for a monthly flow check? | [24 empty] | yes | yes | Yes | yes | yes | yes | yes | yes | yes | yes | yes | Yes | Yes | yes | yes | bar: 2022-08-25 to 2023-04-21
IF SO, what is the 'as found' sampler indicated flow rate (LPM)? | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 |  | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.7 | 16.7 | - | - | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.7 | 17 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7
what is indicated delta cal flow rate? | 16.9 | 16.7 | 16.5 | 16.82 | 16.65 | 16.74 |  | 16.72 | 16.68 | 16.72 | 16.82 | 16.81 | 16.76 | 16.74 | - | 16.51 | 16.66 | - | - | 16.8 | 16.67 | 16.79 | 16.65 | - | 16.52 | 16.79 | 16.66 | 16.71 | 16.67 | 16.76 | 16.3 | 16.4 | 16.2 | 16.5 | 16.6 | 16.51 | 16.8 | 16.65 | 16.83
what is the difference betwen cal and sampler flow (passing < 0.7 LPM) | 0.2 | 0 | 0.2 | 0.12 | 0.05 | 0.04 | 0 | 0.02 | 0.02 | 0.02 | 0.12 | 0.11 | 0.06 | 0.04 | - | 0.19 | 0.04 | - | - | 0.1 | 0.03 | 0.09 | 0.05 | - | 0.18 | 0.21 | 0.04 |  | 0.03 | 0.06 | 0.4 | 0.3 | 0.5 | 0.2 | 0.1 | 0.19 | 0.1 | 0.05 | 0.13
what is the difference betwen cal and design flow (passing < 0.8 LPM) | 0.2 | 0 | 0.2 | 0.12 | 0.05 | 0.04 | 0 | 0.02 | 0.02 | 0.02 | 0.12 | 0.11 | 0.06 | 0.04 | - | 0.19 | 0.04 | - | - | 0.1 | 0.03 | 0.09 | 0.05 | - | 0.28 | 0.09 | 0.04 |  | 0.03 | 0.06 | 0.4 | 0.3 | 0.6 | 0.3 | 0.2 | 0.29 | 0.2 | 0.15 | 0.13
IF ADJUSTMENTS ARE MADE, what is the 'as left' sampler rate (LPM)? | 16.7 | - | - | - | - | - |  | - | - | - | - | 16.8 | - | - | - | - | - | - | - | - | - | - | - | - | - | n/a | n/a | n/a | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.7
what is the 'as left' calibraton rate (LPM) | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 |  | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.7 | 16.7 | - | - | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.6 | 17 | 16.7 |  | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.8 | 16.7 | 16.6
Record the Bam's ambient temp | 23 | 24.7 | 24.7 | 23.9 | 28.7 | 13.9 |  | 9.3 | 7.6 | 20.6 | 5.2 | 6.5 | 6.2 | 5.8 | - | -0.3 | 1.6 | - | - | -2.2 | 2.1 | 5.3 | 24.4 | -
Record the Delta-cal ambient temp (passing = +/-2 deg C) | 24.2 | 23.8 | 24.5 | 22.7 | 28.5 | 10 |  | 5.8 | 8.6 | 21.7 | 3.4 | 7.5 | 5.8 | 5.3 | - | 0.3 | 2 | - | - | -2.5 | 1.8 | 5.5 | 23.8 | -
if adjustment is made, record the new sampler ambient temp | 24.2 | 23.8 | - | - | - | 10 |  | 5.8 | - | - | 3.4 | 7.5 | 5.8 | - | - | - | - | - | - | - | - | - | - | -
Record the Bam's BP | 748 | 149 | 748 | 751 | 743 | 749 |  | 740 | 747 | 754 | 756 | 746 | 742 | 742 | - | 752 | 747 | - | - | 762 | 746 | 752 | 242 | -
Record the Delta-cal BP(passing = +/- 10mmHg) | 747 | 149 | 748 | 751 | 742.5 | 748.5 |  | 739 | 746.5 | 753 | 755.2 | 745.5 | 741.5 | 741.8 | - | 750.7 | 747.7 | - | - | 761.7 | 745.1 | 752.8 | 242.5 | -
if adjustment is made, record the new sampler BP | 747 | - | - | - | - | - |  | - | - | - | - | - | - | - | - | 751 | - | - | - | - | - | - | - | -
Is it time for a monthly leak check? | Y | Y | Y | Y | Y | Y | N | Y | Y | Y | Y | Y | Y | Y | N | Y | Y | N | N | y | Y | Y | Y | N | yes | no | yes | no | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
IF SO, what is the 'as found' leak check value (passing < 0.5 LPM)? | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 |  | 0.4 | 0.5 | 0.4 | 0.5 | 0.5 | 0.4 | 0.4 | - | 0.5 | 0.5 | - | - | 0.5 | 0.5 | 0.5 | 0.4 | - | 0.4 | - | 0.4 | - | 0.3 | 0.3 | 2.8 | 0.1 | 1.3 | 0.4 | 0.3 | 0.3 | 0.4 | 0.3 | 0.4
IF FAILING, what is the 'as left' leak check value? | 0.3 | - | - | - | - | - |  | - | - | - | - | - | - | - | - | - |  | - | - | - | - | - | - | - | 0.5 | n/a | n/a | n/a | 0.3 | - | 0.2 | - | 0.3 | 0.3 | 0.3 | 0.3 | 0.4 | 0.5 | 0.3
Membrane Check
Last m (mg) | 0.794 | 0.798 | 0.794 | 0.797 | 0.801 | 0.801 | 0.804 | 0.804 | 0.799 | 0.803 | 0.801 | 0.805 | 0.801 | 0.803 | 0.801 | 0.802 | 0.8 | 0.797 | 0.797 | 0.799 | 0.799 | NA | 0.803 | 0.801 | 0.786 | 0.78 | 0.781 | 0.781 | 0.782 | 0.78 | 0.782 | 0.783 | 0.783 | 0.783 | 0.785 | 0.785 | 0.79 | 0.788 | 0.786
ABS (mg) | 0.784 | 0.784 | 0.784 | 0.784 | 0.784 | 0.784 | 0.784 | 0.784 | 0.784 | 0.784 | 0.784 | 0.784 | 0.784 | 0.784 | 0.784 | 0.784 | 0.784 | 0.784 | 0.784 | 0.784 | 0.784 | 0.784 | 0.784 | 0.784 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782
Difference (mg) | 0.01 | 0.014 | 0.01 | 0.013 | 0.017 | 0.017 | 0.02 | 0.02 | 0.015 | 0.019 | 0.017 | 0.021 | 0.017 | 0.019 | 0.017 | 0.018 | 0.016 | 0.013 | 0.013 | 0.015 | 0.015 |  | 0.019 | 0.017 | 0.004 | -0.002 | -0.001 | -0.001 | 0 | -0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0.002 | 0 | 0.004
[Last-org] / Org=% difference (<5%) | 1.276% | 1.786% | 1.276% | 1.658% | 2.168% | 2.168% | 2.551% | 2.551% | 1.913% | 2.423% | 2.168% | 2.679% | 2.168% | 2.423% | 2.168% | 2.296% | 2.041% | 1.658% | 1.658% | 1.913% | 1.913% |  | 2.423% | 2.168% | 0% | -0.001% | 0.01% | 0.01% | 0% | 0.02% | 0% | 0% | 0% | 0% | 0% | 0% | 100% | 0% | 0%
Any status error flags showing? If so, list. | [24 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: 2022-08-25 to 2023-04-21
Is date/time correct within 60 seconds? | Y set | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | yes | yes | yes | yes | corrected | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
Are all filter transport systems operational? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Are all tape reel cover knobs secure? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Are hourly data being collected? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Is the PM10 inlet head clean and secure? | OK | OK | OK | CLND | Y | Y | Y | Y | Y | Y | Y | Y | RPLCD | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y
Is the inlet head bug screen clean? | OK | OK | OK | Y | Y | Y | Y | Y | Y | Y | Y | Y | RPLCD | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes, changed | yes | yes | yes
Is the water collection bottle clean? | OK | OK | OK | Y | Y | Y | Y | Y | Y | Y | Y | Y | RPLCD | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes, changed | yes | yes | yes
Are o-rings in good condition? | OK | OK | OK | Y | Y | Y | Y | Y | Y | Y | Y | Y | RPLCD | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
Is there enough filter tape? (what % of the roll remains?) | 50% | 40% | 20% | 100% | 95% | 95% | 80% | 60% | 50% | 20% | 100% | 90% | 60% | 40% | 30% | 100% | 90% | 80% | 60% | 40% | 30% | 95% | 80% | 90% | 50% | yes | yes | yes | yes | yes, changed | 75% | Around 50% | Just over 25% | 100% | 75%% | 60% | 45-50% | 25% | 100%
Does the filter tape show any pinholes or irregular PM deposits? | [24 empty] | no | no | no | no | no | no | pin holes | A few X's | pinhole, top R | small tiny hole started | No | no | no | no | no | bar: OK to Y
Is the nozzle clean? | Y clnd | Y | CLND | Y | CLND | Y | Y | CLND | Y | CLND | Y | Y | CLND | Y | CLND | Y | Y | Y | Y | Y | CLND | Y | Y | Y | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
Is the support vane clean? | Y clnd | Y | CLND | Y | CLND | Y | Y | CLND | Y | CLND | Y | Y | CLND | Y | CLND | Y | Y | Y | Y | Y | CLND | Y | Y | Y | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
Is the capstan clean? | Y clnd | Y | CLND | Y | CLND | Y | Y | CLND | Y | CLND | Y | Y | CLND | Y | CLND | Y | Y | Y | Y | Y | CLND | Y | Y | Y | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
Is the pinch roller clean? | Y clnd | Y | CLND | Y | CLND | Y | Y | CLND | Y | CLND | Y | Y | CLND | Y | CLND | Y | Y | Y | Y | Y | CLND | Y | Y | Y | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
Is the pump muffler clean? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: OK to Y
Is anything touching the (HOT) pump? | [24 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: OK to Y
Is the temperature probe clean/tight? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: OK to Y
Is the temperature probe signal connection tight? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: OK to Y
Is the signal connections on the rear panel tight? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: OK to Y
Is the AC air filter clean? | rotated | Y | Y | Y | - | Y | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | yes | - | - | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | changed
Is the evap. coil (base of AC unit) clean? | clnd | Y | Y | Y | - | CLND | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | yes | - | - | - | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | cleaned
Is the AC comp. drain line clean? | [24 empty] | yes | - | - | - | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: clnd to -
Is the AC temp. setting ~82F? | 82 |  | Y |  | - | Y | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | yes | - | - | - | yes | yes | yes | yes | yes | yes | - | yes | yes | yes | yes
General Info
Is the site secure with no signs of damage? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: clnd to -
Are all monitors in SAMPLE mode (no standby)? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: clnd to -
Is communication system operational? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: clnd to -
Is climate control working? Note current inside temp in Deg F | Y | Y | 86.19 | 84.52 | 31.19 | 24.72 | 22.36 | 21.29 | 22.39 | 28.72 | 19.47 | 19.65 | 21.5 | 18.5 | 21.4 | 13.99 | 18.92 | - | 23.6 | 18.12 | 18.5 | 19.78 | 30.67 | 21.28 | yes | - | yes | yes | 82F | yes | 82F | 83 | 83 | yes | yes | yes | yes | yes | yes
*Are there any unusual site noises? | [24 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: clnd to -
*Are there any unusual site observations? | [24 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: clnd to -
Is the site clean and manicured? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes, no fence | yes | yes | yes | yes | yes | tall grass | bar: clnd to -
Are there any site safety issues? | [24 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: clnd to -
*Are supplies and expendables needed or ordered? | [24 empty] | ok | ok | yes | yes | no | no | no | no | no | Long Q-tips | no | no | no | no | no | bar: clnd to -
Operator Initials: | [24 empty] | JCG | jcg | gh | jcg | ds | jcg | ds | JCG | JCG | JCG | JCG | JCG | JCG | JCG | DS | bar: clnd to -
Departure Time: | 12:30 | 8:35 | 10:10 | 9:45 | 9:40 | 10:40 | 9:15 | 9:50 | 11:45 | 12:20 | 11:20 | 9:20 | 11:20 | 11:30 | 12:10 | 11:30 | 10:20 | 9:05 | 11:10 | 10:10 | 11:10 | 11:50 | 11:10 | 10:50 |  | 1136 | 1201 | 1145 | 955 | 1155 | 15:15 | 1315 | 1340 | 1522 | 1405 | 1330 | 1210 | 1150 | 15:15
*If Yes call Dave Szatkowski to discuss (512)653-1737

## ST-CE
KCBX ST-CE Operator Checklist |  | 10 weeks
Arrival Date: (to insert new dates selet Column C. - Right click - insert 1 left) | 2022-08-25 | 2022-09-02 | 2022-09-08 | 2022-09-16 | 2022-09-21 | 2022-09-27 | 2022-10-07 | 2022-10-14 | 2022-10-19 | 2022-11-02 | 2022-11-14 | 2022-11-28 | 2022-12-13 | 2022-12-28 | 2023-01-05 | 2023-01-13 | 2023-01-20 | 2023-01-27 | 2023-02-06 | 2023-02-24 | 2023-03-10 | 2023-03-30 | 2023-04-14 | 2023-04-21 | 2015-10-28 | 2015-04-06 | 2015-04-14 | 2015-04-21 | 2015-05-06 | 2015-05-20 | 2015-06-03 | 2015-06-19 | 2015-07-06 | 2015-07-21 | 2015-08-04 | 2015-08-17 | 2015-09-03 | 2015-09-16 | 2015-10-01
Arrival Time: | 16:40 | 8:50 | 10:20 | 10:05 | 10:00 | 11:00 | 9:20 | 10:05 | 12:00 | 12:25 | 12:25 | 9:20 | 11:20 | 11:35 | 1:10 | 11:35 | 10:20 | 9:05 | 11:10 | 10:10 | 11:10 | 11:50 | 11:10 | 10:50 | 1111 |  |  | 1125 | 9:05 | 11:14 | 14:00 | 1245 | 1303 | 1427 | 1320 | 1255 | 1136 | 1058 | 13:40
BAM-1020 PM10 Monitor
Is it time for a monthly flow check? | [24 empty] | yes | yes | Yes | yes | yes | yes | yes | yes | yes | yes | yes | Yes | Yes | yes | yes | bar: 2022-08-25 to 2023-04-21
IF SO, what is the 'as found' sampler indicated flow rate (LPM)? | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 |  | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.7 | 16.7 | - | - | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.7 | 17 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7
what is indicated delta cal flow rate? | 16.9 | 16.72 | 16.8 | 16.67 | 16.76 | 16.64 |  | 16.46 | 16.38 | 16.65 | 16.49 | 16.55 | 16.49 | 16.78 | - | 16.74 | 16.81 | - | - | 16.84 | 16.73 | 16.82 | 16.92 | - | 16.52 | 16.79 | 16.66 | 16.71 | 16.67 | 16.76 | 16.3 | 16.4 | 16.2 | 16.5 | 16.6 | 16.51 | 16.8 | 16.65 | 16.83
what is the difference betwen cal and sampler flow (passing < 0.7 LPM) | 0.2 | 0.02 | 0.1 | 0.03 | 0.06 | 0.06 | 0 | 0.24 | 0.32 | 0.05 | 0.21 | 0.15 | 0.21 | 0.08 | - | 0.04 | 0.11 | - | - | 0.14 | 0.03 | 0.12 | 0.22 | - | 0.18 | 0.21 | 0.04 |  | 0.03 | 0.06 | 0.4 | 0.3 | 0.5 | 0.2 | 0.1 | 0.19 | 0.1 | 0.05 | 0.13
what is the difference betwen cal and design flow (passing < 0.8 LPM) | 0.2 | 0.02 | 0.1 | 0.03 | 0.06 | 0.06 | 0 | 0.24 | 0.32 | 0.05 | 0.21 | 0.15 | 0.21 | 0.08 | - | 0.04 | 0.11 | - | - | 0.14 | 0.03 | 0.12 | 0.22 | - | 0.28 | 0.09 | 0.04 |  | 0.03 | 0.06 | 0.4 | 0.3 | 0.6 | 0.3 | 0.2 | 0.29 | 0.2 | 0.15 | 0.13
IF ADJUSTMENTS ARE MADE, what is the 'as left' sampler rate (LPM)? | 16.7 | - | - | - | - | - |  | - | - | - | 16.5 | 16.6 | 16.5 | - | - | - | - | - | - | - | - | - | 16.9 | - | - | n/a | n/a | n/a | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.7
what is the 'as left' calibraton rate (LPM) | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 |  | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.7 | 16.7 | - | - | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.6 | 17 | 16.7 |  | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.8 | 16.7 | 16.6
Record the Bam's ambient temp | 25 | 25.9 | 26.5 | 26.7 | 31.2 | 15.7 |  | 10.6 | 9.1 | 21.6 | 6.3 | 7.6 | 7.2 | 6.5 | - | 1.2 | 1.3 | - | - | -2.9 | 1.9 | 5.6 | 24.8 | -
Record the Delta-cal ambient temp (passing = +/-2 deg C) | 26.2 | 25.2 | 27.6 | 25.7 | 32.7 | 11.6 |  | 4.6 | 8.2 | 22.2 | 5.4 | 7.6 | 7.1 | 6.8 | - | 0 | 1.5 | - | - | -3.1 | 1.9 | 5.3 | 25.5 | -
if adjustment is made, record the new sampler ambient temp | 26.2 | - | 27.6 | - | - | 11.6 |  | - | - | - | 5.4 | - | - | - | - | 0 | - | - | - | - | - | - | - | -
Record the Bam's BP | 744 | 750 | 749 | 752 | 744 | 750 |  | 740 | 747 | 754 | 756 | 747 | 742 | 742 | - | 751 | 748 | - | - | 762 | 746 | 753 | 742 | -
Record the Delta-cal BP(passing = +/- 10mmHg) | 746 | 748.5 | 747.5 | 751 | 742.5 | 748.5 |  | 739 | 746.5 | 753 | 754.8 | 745.5 | 741 | 741.8 | - | 750.2 | 747.6 | - | - | 761 | 745 | 753 | 742.5 | -
if adjustment is made, record the new sampler BP | 746 | - | - | - | - | - |  | - | - | - | - | 746 | 741 | - | - | - | - | - | - | - | - | - | - | -
Is it time for a monthly leak check? | Y | Y | Y | Y | Y | Y | N | Y | Y | Y | Y | y | Y | Y | N | Y | Y | N | N | y | Y | Y | Y | N | yes | no | yes | no | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
IF SO, what is the 'as found' leak check value (passing < 0.5 LPM)? | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 |  | 0.3 | 0.2 | 0.2 | 0.3 | 0.3 | 0.2 | 0.2 | - | 0.3 | 0.4 | - | - | 0.3 | 0.2 | 0.3 | 0.2 | - | 0.4 | - | 0.4 | - | 0.3 | 0.3 | 2.8 | 0.1 | 1.3 | 0.4 | 0.3 | 0.3 | 0.4 | 0.3 | 0.4
IF FAILING, what is the 'as left' leak check value? | 0.2 | - | - | - | - | - |  |  |  |  | - | - | - | - | - | - |  | - | - | - | - | - | - | - | 0.5 | n/a | n/a | n/a | 0.3 | - | 0.2 | - | 0.3 | 0.3 | 0.3 | 0.3 | 0.4 | 0.5 | 0.3
Membrane Check
Last m (mg) |  | 0.783 | 0.786 | 0.786 | 0.782 | 0.784 | 0.785 | 0.783 | 0.784 | 0.783 | 0.781 | 0.782 | 0.794 | 0.786 | 0.789 | 0.791 | 0.789 | 0.788 | 0.787 | 0.786 | 0.785 | NA | 0.788 | 0.785 | 0.786 | 0.78 | 0.781 | 0.781 | 0.782 | 0.78 | 0.782 | 0.783 | 0.783 | 0.783 | 0.785 | 0.785 | 0.79 | 0.788 | 0.786
ABS (mg) |  | 0.778 | 0.778 | 0.778 | 0.778 | 0.778 | 0.778 | 0.778 | 0.778 | 0.778 | 0.778 | 0.778 | 0.778 | 0.778 | 0.778 | 0.778 | 0.778 | 0.778 | 0.778 | 0.778 | 0.778 | 0.778 | 0.778 | 0.778 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782
Difference (mg) |  | 0.005 | 0.008 | 0.008 | 0.004 | 0.006 | 0.007 | 0.005 | 0.006 | 0.005 | 0.003 | 0.004 | 0.016 | 0.008 | 0.011 | 0.013 | 0.011 | 0.01 | 0.009 | 0.008 | 0.007 |  | 0.01 | 0.007 | 0.004 | -0.002 | -0.001 | -0.001 | 0 | -0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0.002 | 0 | 0.004
[Last-org] / Org=% difference (<5%) |  | 0.6427% | 1.028% | 1.028% | 0.5141% | 0.7712% | 0.8997% | 0.6427% | 0.7712% | 0.6427% | 0.3856% | 0.5141% | 2.057% | 1.028% | 1.414% | 1.671% | 1.414% | 1.285% | 1.157% | 1.028% | 0.8997% |  | 1.285% | 0.8997% | 0% | -0.001% | 0.01% | 0.01% | 0% | 0.02% | 0% | 0% | 0% | 0% | 0% | 0% | 100% | 0% | 0%
Any status error flags showing? If so, list. | [24 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: 2022-08-25 to 2023-04-21
Is date/time correct within 60 seconds? | Y set | Y | Y | Y | Y | Y | Y | Y | ADJ | Y | Y | y | ADJ | Y | Y | Y | Y | Y | ADJ | Y | Y | Y | Y | Y | yes | yes | yes | yes | corrected | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
Are all filter transport systems operational? | OK | OK | OK | Y | Y | Y | Y | Y | Y | Y | Y | y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
Are all tape reel cover knobs secure? | OK | OK | OK | Y | Y | Y | Y | Y | Y | Y | Y | y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
Are hourly data being collected? | OK | OK | OK | Y | Y | Y | Y | Y | Y | Y | Y | y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
Is the PM10 inlet head clean and secure? | OK | OK | OK | CLND | Y | Y | Y | Y | Y | Y | Y | y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y
Is the inlet head bug screen clean? | OK | OK | OK | Y | Y | Y | Y | Y | Y | Y | Y | y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes, changed | yes | yes | yes
Is the water collection bottle clean? | OK | OK | OK | Y | Y | Y | Y | Y | Y | Y | Y | y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes, changed | yes | yes | yes
Are o-rings in good condition? | OK | OK | OK | Y | Y | Y | Y | Y | Y | Y | Y | y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
Is there enough filter tape? (what % of the roll remains?) | 50% | 60% | 30% | 20% | 100% | 95% | 90% | 80% | 80% | 50% | 30% | 100% | 80%% | 60% | 50% | 40% | 20% | 100% | 90% | 50%% | 40% | 20% | 90% | 90% | 50% | yes | yes | yes | yes | yes, changed | 75% | Around 50% | Just over 25% | 100% | 75%% | 60% | 45-50% | 25% | 100%
Does the filter tape show any pinholes or irregular PM deposits? | NO | N | cut outs | N | N | N | N | N | N | N | N | n | N | N | N | N | N | N | N | N | N | N | N | N | no | no | no | no | no | no | pin holes | A few X's | pinhole, top R | small tiny hole started | No | no | no | no | no
Is the nozzle clean? | Y clnd | Y | CLND | Y | Y | CLND | Y | Y | CLND | Y | CLND | y | CLND | Y | Y | Y | CLND | Y | Y | Y | CLND | Y | Y | Y | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
Is the support vane clean? | Y clnd | Y | CLND | Y | Y | CLND | Y | Y | CLND | Y | CLND | y | CLND | Y | Y | Y | CLND | Y | Y | Y | CLND | Y | Y | Y | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
Is the capstan clean? | Y clnd | Y | CLND | Y | Y | CLND | Y | Y | CLND | Y | CLND | y | CLND | Y | Y | Y | CLND | Y | Y | Y | CLND | Y | Y | Y | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
Is the pinch roller clean? | Y clnd | Y | CLND | Y | Y | CLND | Y | Y | CLND | Y | CLND | y | CLND | Y | Y | Y | CLND | Y | Y | Y | CLND | Y | Y | Y | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
Is the pump muffler clean? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: OK to Y
Is anything touching the (HOT) pump? | [24 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: OK to Y
Is the temperature probe clean/tight? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: OK to Y
Is the temperature probe signal connection tight? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: OK to Y
Is the signal connections on the rear panel tight? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: OK to Y
Is the AC air filter clean? | rotated | Y | Y | Y | - | Y | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | yes | - | - | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | changed
Is the evap. coil (base of AC unit) clean? | clnd | Y | Y | Y | - | CLND | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | yes | - | - | - | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | cleaned
Is the AC comp. drain line clean? | [24 empty] | yes | - | - | - | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: clnd to -
Is the AC temp. setting ~82F? | 80 |  | Y | - | - | Y | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | yes | - | - | - | yes | yes | yes | yes | yes | yes | - | yes | yes | yes | yes
General Info
Is the site secure with no signs of damage? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: clnd to -
Are all monitors in SAMPLE mode (no standby)? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: clnd to -
Is communication system operational? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: clnd to -
Is climate control working? Note current inside temp in Deg F | Y | Y | 87.89 | 91.17 | 33.92 | 30.98 | 24.42 | 24.68 | 26.28 | 33.7 | 23.72 | 29.29 | 27.5 | 27.9 | 27.37 | 27.34 | 27.56 | 20.69 | 28.62 | 26.62 | 26.6 | 21.3 | 32.37 | 28.23 | yes | - | yes | yes | 82F | yes | 82F | 83 | 83 | yes | yes | yes | yes | yes | yes
*Are there any unusual site noises? | [24 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: clnd to -
*Are there any unusual site observations? | [24 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: clnd to -
Is the site clean and manicured? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes, no fence | yes | yes | yes | yes | yes | tall grass | bar: clnd to -
Are there any site safety issues? | [24 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: clnd to -
*Are supplies and expendables needed or ordered? | [24 empty] | ok | ok | yes | yes | no | no | no | no | no | Long Q-tips | no | no | no | no | no | bar: clnd to -
Operator Initials: | [24 empty] | JCG | jcg | gh | jcg | ds | jcg | ds | JCG | JCG | JCG | JCG | JCG | JCG | JCG | DS | bar: clnd to -
Departure Time: | 17:20 | 9:05 | 10:50 | 10:20 | 10:15 | 11:20 | 9:35 | 10:15 | 12:15 | 12:35 | 11:40 | 9:35 | 11:30 | 11:40 | 1:20 | 11:45 | 10:40 | 9:15 | 11:20 | 10:20 | 11:20 | 12:00 | 11:20 | 11:00 |  | 1136 | 1201 | 1145 | 955 | 1155 | 15:15 | 1315 | 1340 | 1522 | 1405 | 1330 | 1210 | 1150 | 15:15
*If Yes call Dave Szatkowski to discuss (512)653-1737

## ST-SW
KCBX ST-SW Operator Checklist |  | 10 weeks
Arrival Date: (to insert new dates selet Column C. - Right click - insert 1 left) | 2022-08-25 | 2022-09-02 | 2022-09-08 | 2022-09-16 | 2022-09-21 | 2022-09-27 | 2022-10-07 | 2022-10-14 | 2022-10-19 | 2022-11-02 | 2022-11-14 | 2022-11-28 | 2022-12-13 | 2022-12-28 | 2023-01-05 | 2023-01-13 | 2023-01-20 | 2023-01-27 | 2023-02-06 | 2023-02-24 | 2023-03-10 | 2023-03-30 | 2023-04-14 | 2023-04-21 | 2015-10-28 | 2015-04-06 | 2015-04-14 | 2015-04-21 | 2015-05-06 | 2015-05-20 | 2015-06-03 | 2015-06-19 | 2015-07-06 | 2015-07-21 | 2015-08-04 | 2015-08-17 | 2015-09-03 | 2015-09-16 | 2015-10-01
Arrival Time: | 15:40 | 12:45 | 10:50 | 10:25 | 10:20 | 11:30 | 9:40 | 10:20 | 12:20 | 12:40 | 12:00 | 9:40 | 12:30 | 11:45 | 12:15 | 11:45 | 10:40 | 9:15 | 11:20 | 10:20 | 10:00 | 12:00 | 11:20 | 11:00 | 1111 |  |  | 1125 | 9:05 | 11:14 | 14:00 | 1245 | 1303 | 1427 | 1320 | 1255 | 1136 | 1058 | 13:40
BAM-1020 PM10 Monitor
Is it time for a monthly flow check? | [24 empty] | yes | yes | Yes | yes | yes | yes | yes | yes | yes | yes | yes | Yes | Yes | yes | yes | bar: 2022-08-25 to 2023-04-21
IF SO, what is the 'as found' sampler indicated flow rate (LPM)? | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 |  | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.7 | 16.7 | - | - | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.7 | 17 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7
what is indicated delta cal flow rate? | 17.1 | 16.85 | 16.86 | 16.89 | 16.88 | 16.97 |  | 16.9 | 16.83 | 16.86 | 16.97 | 16.83 | 16.71 | 16.72 | - | 16.68 | 16.61 | - | - | 16.71 | 16.78 | 16.9 | 17.18 | - | 16.52 | 16.79 | 16.66 | 16.71 | 16.67 | 16.76 | 16.3 | 16.4 | 16.2 | 16.5 | 16.6 | 16.51 | 16.8 | 16.65 | 16.83
what is the difference betwen cal and sampler flow (passing < 0.7 LPM) | 0.4 | 0.15 | 0.16 | 0.19 | 0.18 | 0.27 | 0 | 0.2 | 0.13 | 0.16 | 0.27 | 0.13 | 0.01 | 0.02 | - | 0.02 | 0.09 | - | - | 0.01 | 0.08 | 0.2 | 0.48 | - | 0.18 | 0.21 | 0.04 |  | 0.03 | 0.06 | 0.4 | 0.3 | 0.5 | 0.2 | 0.1 | 0.19 | 0.1 | 0.05 | 0.13
what is the difference betwen cal and design flow (passing < 0.8 LPM) | 0.4 | 0.15 | 0.16 | 0.19 | 0.18 | 0.27 | 0 | 0.2 | 0.13 | 0.16 | 0.27 | 0.13 | 0.01 | 0.02 | - | 0.02 | 0.09 | - | - | 0.01 | 0.08 | 0.2 | 0.48 | - | 0.28 | 0.09 | 0.04 |  | 0.03 | 0.06 | 0.4 | 0.3 | 0.6 | 0.3 | 0.2 | 0.29 | 0.2 | 0.15 | 0.13
IF ADJUSTMENTS ARE MADE, what is the 'as left' sampler rate (LPM)? | 16.7 | 16.8 | - | - | - | - |  | 16.9 | - | - | 17 | 16.8 | - | - | - | - | - | - | - | - | - | 16.9 | 17.1 | - | - | n/a | n/a | n/a | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.7
what is the 'as left' calibraton rate (LPM) | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 |  | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.7 | 16.7 | - | - | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.6 | 17 | 16.7 |  | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.8 | 16.7 | 16.6
Record the Bam's ambient temp | 25.1 | 31.1 | 27.6 | 25.5 | 30.9 | 15.7 |  | 10.4 | 9 | 21.4 | 7.8 | 7.6 | 7.6 | 5.8 | - | 0.5 | 1.9 | - | - | -2.8 | 2 | 7.5 | 25.4 | -
Record the Delta-cal ambient temp (passing = +/-2 deg C) | 26 | 32.7 | 30.7 | 25.5 | 33.7 | 11.7 |  | 4.7 | 9.5 | 21.6 | 5.1 | 7.6 | 7.3 | 5.9 | - | 0.3 | 1.7 | - | - | -2.5 | 1.2 | 5.8 | 24.8 | -
if adjustment is made, record the new sampler ambient temp | 26 | 32.7 | 30.7 | - | 33.7 | 11.7 |  | - | - | - | 5.1 | - | 7.3 | - | - | - | - | - | - | - | 1.2 | - | - | -
Record the Bam's BP | 745 | 749 | 748 | 752 | 744 | 753 |  | 741 | 751 | 754 | 755 | 748 | 743 | 742 | - | 755 | 743 | - | - | 763 | 745 | 754 | 745 | -
Record the Delta-cal BP(passing = +/- 10mmHg) | 746 | 749 | 747.5 | 751 | 743.5 | 748.5 |  | 739 | 747 | 753 | 754.5 | 746 | 741 | 741.7 | - | 750.2 | 747.4 | - | - | 761.2 | 745 | 753 | 742.5 | -
if adjustment is made, record the new sampler BP | 746 | - | - | - | - | 748.5 |  | - | - | - | - | 746 | 741 | - | - | 750 | 747 | - | - | 762 | - | - | 742 | -
Is it time for a monthly leak check? | Y | Y | Y | Y | Y | Y | N | Y | Y | Y | Y | 7 | Y | Y | N | Y | Y | N | N | Y | Y | Y | Y | N | yes | no | yes | no | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
IF SO, what is the 'as found' leak check value (passing < 0.5 LPM)? | 0.2 | 0.3 | 0.3 | 0.3 | 0.2 | 0.3 |  | 0.3 | 0.3 | 0.3 | 0.3 | 0.4 | 0.4 | 0.4 | - | 0.4 | 0.3 | - | - | 0.5 | 0.4 | 0.4 | 0.4 | - | 0.4 | - | 0.4 | - | 0.3 | 0.3 | 2.8 | 0.1 | 1.3 | 0.4 | 0.3 | 0.3 | 0.4 | 0.3 | 0.4
IF FAILING, what is the 'as left' leak check value? | 0.2 | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - |  | - | - | - | - | - | - | - | 0.5 | n/a | n/a | n/a | 0.3 | - | 0.2 | - | 0.3 | 0.3 | 0.3 | 0.3 | 0.4 | 0.5 | 0.3
Membrane Check
Last m (mg) |  | 0.802 | 0.801 | 0.8 | 0.801 | 0.803 | 0.803 | 0.805 | 0.801 | 0.802 | 0.804 | 0.807 | 0.806 | 0.808 | 0.811 | 0.81 | 0.806 | 0.811 | 0.81 | 0.809 | 0.812 | NA | 0.812 | 0.809 | 0.786 | 0.78 | 0.781 | 0.781 | 0.782 | 0.78 | 0.782 | 0.783 | 0.783 | 0.783 | 0.785 | 0.785 | 0.79 | 0.788 | 0.786
ABS (mg) |  |  | 0.802 | 0.802 | 0.802 | 0.802 | 0.802 | 0.802 | 0.802 | 0.802 | 0.802 | 0.802 | 0.802 | 0.802 | 0.802 | 0.802 | 0.802 | 0.802 | 0.802 | 0.802 | 0.802 | 0.802 | 0.802 | 0.802 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782
Difference (mg) |  |  | 0.001 | 0.002 | 0.001 | 0.001 | 0.001 | 0.003 | 0.001 | 0 | 0.002 | 0.005 | 0.004 | 0.006 | 0.009 | 0.008 | 0.004 | 0.009 | 0.008 | 0.007 | 0.01 |  | 0.01 | 0.007 | 0.004 | -0.002 | -0.001 | -0.001 | 0 | -0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0.002 | 0 | 0.004
[Last-org] / Org=% difference (<5%) |  |  | -0.1247% | -0.2494% | -0.1247% | 0.1247% | 0.1247% | 0.3741% | -0.1247% | 0% | 0.2494% | 0.6234% | 0.4988% | 0.7481% | 1.122% | 0.9975% | 0.4988% | 1.122% | 0.9975% | 0.8728% | 1.247% |  | 1.247% | 0.8728% | 0% | -0.001% | 0.01% | 0.01% | 0% | 0.02% | 0% | 0% | 0% | 0% | 0% | 0% | 100% | 0% | 0%
Any status error flags showing? If so, list. | [24 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: 2022-08-25 to 2023-04-21
Is date/time correct within 60 seconds? | Y set | Y | Y | Y | Y | Y | Y | Y | Y | Y | ADJ | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | yes | yes | yes | yes | corrected | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
Are all filter transport systems operational? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Are all tape reel cover knobs secure? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Are hourly data being collected? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Is the PM10 inlet head clean and secure? | OK | OK | OK | CLND | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y
Is the inlet head bug screen clean? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes, changed | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Is the water collection bottle clean? | OK | clnd | CLND | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes, changed | yes | yes | yes
Are o-rings in good condition? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: OK to Y
Is there enough filter tape? (what % of the roll remains?) | 55% | 50 | 40% | 30% | 25% | 15% | 100% | 95% | 90% | 80% | 50% | 30% | 100% | 90%% | 60% | 50% | 40% | 30% | 100% | 80% | 60% | 30% | 100% | 70% | 50% | yes | yes | yes | yes | yes, changed | 75% | Around 50% | Just over 25% | 100% | 75%% | 60% | 45-50% | 25% | 100%
Does the filter tape show any pinholes or irregular PM deposits? | [24 empty] | no | no | no | no | no | no | pin holes | A few X's | pinhole, top R | small tiny hole started | No | no | no | no | no | bar: OK to Y
Is the nozzle clean? | Y clnd | Y | CLND | Y | CLND | Y | Y | CLND | Y | CLND | Y | Y | CLND | Y | Y | Y | CLND | Y | Y | Y | CLND | Y | Y | Y | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
Is the support vane clean? | Y clnd | Y | CLND | Y | CLND | Y | Y | CLND | Y | CLND | Y | Y | CLND | Y | Y | Y | CLND | Y | Y | Y | CLND | Y | Y | Y | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
Is the capstan clean? | Y clnd | Y | CLND | Y | CLND | Y | Y | CLND | Y | CLND | Y | Y | CLND | Y | Y | Y | CLND | Y | Y | Y | CLND | Y | Y | Y | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
Is the pinch roller clean? | Y clnd | Y | CLND | Y | CLND | Y | Y | CLND | Y | CLND | Y | Y | CLND | Y | Y | Y | CLND | Y | Y | Y | CLND | Y | Y | Y | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
Is the pump muffler clean? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: OK to Y
Is anything touching the (HOT) pump? | [24 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: OK to Y
Is the temperature probe clean/tight? | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | CLND | Y | Y | Y | Y | Y | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
Is the temperature probe signal connection tight? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: OK to Y
Is the signal connections on the rear panel tight? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: OK to Y
Is the AC air filter clean? | clnd | Y | Y | Y | Y | RPLCD | - | - | - | Y | - | - | - | - | - | - | - | - | - | - | - | - | - | - | yes | - | - | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | changed
Is the evap. coil (base of AC unit) clean? | clnd | Y | Y | Y | Y | CLND | - | - | - | CLND | - | - | - | - | - | - | - | - | - | - | - | - | - | - | yes | - | - | - | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | cleaned
Is the AC comp. drain line clean? | [24 empty] | yes | - | - | - | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: clnd clnd to - -
Is the AC temp. setting ~82F? | 78 |  | Y | - | Y | Y | - | - | - | Y | - | - | - | - | - | - | - | - | - | - | - | - | - | - | yes | - | - | - | yes | yes | yes | yes | yes | yes | - | yes | yes | yes | yes
General Info
Is the site secure with no signs of damage? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: clnd clnd to - -
Are all monitors in SAMPLE mode (no standby)? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: clnd clnd to - -
Is communication system operational? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: clnd clnd to - -
Is climate control working? Note current inside temp in Deg F | Y | Y | 84.58 | 96.8 | 27.64 | 31.76 | 29.11 | 28.69 | 24.65 | 25.25 | 32.04 | 27.03 | 25.7 | 24.5 | 20.69 | 17.83 | 18.56 | - | 25.58 | 16.94 | 18.9 | 27.8 | 34.96 | 27.97 | yes | - | yes | yes | 82F | yes | 82F | 83 | 83 | yes | yes | yes | yes | yes | yes
*Are there any unusual site noises? | [24 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: clnd clnd to - -
*Are there any unusual site observations? | [24 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: clnd clnd to - -
Is the site clean and manicured? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes, no fence | yes | yes | yes | yes | yes | tall grass | bar: clnd clnd to - -
Are there any site safety issues? | [24 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: clnd clnd to - -
*Are supplies and expendables needed or ordered? | [24 empty] | ok | ok | yes | yes | no | no | no | no | no | Long Q-tips | no | no | no | no | no | bar: clnd clnd to - -
Operator Initials: | [24 empty] | JCG | jcg | gh | jcg | ds | jcg | ds | JCG | JCG | JCG | JCG | JCG | JCG | JCG | DS | bar: clnd clnd to - -
Departure Time: | 16:30 | 1:00 | 11:10 | 10:35 | 11:30 | 11:50 | 10:15 | 10:35 | 12:40 | 2:00 | 12:15 | 10:00 | 12:45 | 12:00 | 12:30 | 11:55 | 10:55 | 9:25 | 11:45 | 10:35 | 10:10 | 12:10 | 11:35 | 11:10 |  | 1136 | 1201 | 1145 | 955 | 1155 | 15:15 | 1315 | 1340 | 1522 | 1405 | 1330 | 1210 | 1150 | 15:15
*If Yes call Dave Szatkowski to discuss (512)653-1737

## ST-NW
KCBX ST-NW Operator Checklist |  | 10 weeks
Arrival Date: (to insert new dates selet Column C. - Right click - insert 1 left) | 2022-08-25 | 2022-09-02 | 2022-09-08 | 2022-09-16 | 2022-09-21 | 2022-09-27 | 2022-10-07 | 2022-10-14 | 2022-10-19 | 2022-11-02 | 2022-11-14 | 2022-11-28 | 2022-12-13 | 2022-12-28 | 2023-01-05 | 2023-01-13 | 2023-01-20 | 2023-01-27 | 2023-02-06 | 2023-02-24 | 2023-03-10 | 2023-03-30 | 2023-04-14 | 2023-04-21 | / | 2015-10-28 | 2015-04-06 | 2015-04-14 | 2015-04-21 | 2015-05-06 | 2015-05-20 | 2015-06-03 | 2015-06-19 | 2015-07-06 | 2015-07-21 | 2015-08-04 | 2015-08-17 | 2015-09-03 | 2015-09-16 | 2015-10-01
Arrival Time: | 14:30 | 10:05 | 11:15 | 11:00 | 11:35 | 12:00 | 10:20 | 10:40 | 10:20 | 2:00 | 12:15 | 10:05 | 11:40 | 12:00 | 12:30 | 12:00 | 11:15 | 9:30 | 12:00 | 10:45 | 10:10 | 12:00 | 11:35 | 11:10 |  | 1111 |  |  | 1125 | 9:05 | 11:14 | 14:00 | 1245 | 1303 | 1427 | 1320 | 1255 | 1136 | 1058 | 13:40
BAM-1020 PM10 Monitor
Is it time for a monthly flow check? | [25 empty] | yes | yes | Yes | yes | yes | yes | yes | yes | yes | yes | yes | Yes | Yes | yes | yes | bar: 2022-08-25 to 2023-04-21
IF SO, what is the 'as found' sampler indicated flow rate (LPM)? | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 |  | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.7 | 16.7 | - | - | 16.7 | 16.7 | 16.7 | 16.7 | - |  | 16.7 | 17 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7
what is indicated delta cal flow rate? | 16.9 | 16.82 | 16.96 | 16.84 | 16.87 | 16.79 |  | 16.57 | 16.78 | 16.75 | 16.61 | 16.65 | 16.56 | 16.79 | - | 16.67 | 16.8 | - | - | 16.79 | 16.82 | 16.92 | 16.62 | - |  | 16.52 | 16.79 | 16.66 | 16.71 | 16.67 | 16.76 | 16.3 | 16.4 | 16.2 | 16.5 | 16.6 | 16.51 | 16.8 | 16.65 | 16.83
what is the difference betwen cal and sampler flow (passing < 0.7 LPM) | 0.2 | 0.12 | 0.26 | 0.14 | 0.17 | 0.09 | 0 | 0.13 | 0.08 | 0.05 | 0.09 | 0.05 | 0.14 | 0.09 | - | 0.03 | 0.1 | - | - | 0.09 | 0.12 | 0.22 | 0.08 | - |  | 0.18 | 0.21 | 0.04 |  | 0.03 | 0.06 | 0.4 | 0.3 | 0.5 | 0.2 | 0.1 | 0.19 | 0.1 | 0.05 | 0.13
what is the difference betwen cal and design flow (passing < 0.8 LPM) | 0.2 | 0.12 | 0.26 | 0.14 | 0.17 | 0.09 | 0 | 0.13 | 0.08 | 0.05 | 0.09 | 0.05 | 0.14 | 0.09 | - | 0.03 | 0.1 | - | - | 0.09 | 0.12 | 0.22 | 0.08 | - |  | 0.28 | 0.09 | 0.04 |  | 0.03 | 0.06 | 0.4 | 0.3 | 0.6 | 0.3 | 0.2 | 0.29 | 0.2 | 0.15 | 0.13
IF ADJUSTMENTS ARE MADE, what is the 'as left' sampler rate (LPM)? | 16.7 | 16.8 | - | - | - | - |  | - | - | - | - | - | 16.6 | - | - | - | - | - | - | - | - | 16.9 | - | - |  | - | n/a | n/a | n/a | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.7
what is the 'as left' calibraton rate (LPM) | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 |  | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.7 | 16.7 | - | - | 16.7 | 16.7 | 16.7 | 16.7 | - |  | 16.6 | 17 | 16.7 |  | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.8 | 16.7 | 16.6
Record the Bam's ambient temp | 24 | 27.2 | 27.6 | 26.2 | 31.6 | 14.2 |  | 10.9 | 6.6 | 22.4 | 6.7 | 8.2 | 7 | 6.2 | - | 0.8 | 1.7 | - | - | -2.5 | 2.1 | 6.7 | 23.8 | -
Record the Delta-cal ambient temp (passing = +/-2 deg C) | 25.4 | 27.7 | 33.7 | 27.2 | 34.4 | 9.5 |  | 5.2 | 7 | 22.7 | 6.3 | 8.6 | 6.8 | 6.5 | - | 0.5 | 1.7 | - | - | -2.9 | 1.6 | 6.3 | 23.7 | -
if adjustment is made, record the new sampler ambient temp | 25.4 | 27.7 | 33.7 | - | 34.4 | 9.5 |  | - | - | - | 6.3 | 8.6 | - | - | - | - | - | - | - | - | 1.6 | - | - | -
Record the Bam's BP | 744 | 749 | 747 | 751 | 743 | 749 |  | 739 | 747 | 752 | 755 | 746 | 741 | 742 | - | 751 | 748 | - | - | 762 | 746 | 753 | 242 | -
Record the Delta-cal BP(passing = +/- 10mmHg) | 746 | 749 | 747.5 | 750.5 | 743.5 | 749 |  | 738.5 | 747 | 752 | 754.4 | 746 | 741 | 741.8 | - | 750.2 | 747.7 | - | - | 761 | 745 | 753 | 242.5 | -
if adjustment is made, record the new sampler BP | 746 | - | - | - | - | - |  | - | - | - | - | - | - | - | - | - | 747 | - | - | - | - | - | - | -
Is it time for a monthly leak check? | [25 empty] | yes | no | yes | no | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
IF SO, what is the 'as found' leak check value (passing < 0.5 LPM)? | 0.2 | 0.3 | 0.2 | 0.2 | 0.2 | 0.2 |  | 0.5 | 0.4 | 0.4 | 0.4 | 0.4 | 0.3 | 0.3 | - | 0.4 | 0.4 | - | - | 0.5 | 0.3 | 0.5 | 0.4 | - |  | 0.4 | - | 0.4 | - | 0.3 | 0.3 | 2.8 | 0.1 | 1.3 | 0.4 | 0.3 | 0.3 | 0.4 | 0.3 | 0.4
IF FAILING, what is the 'as left' leak check value? | 0.2 | - | - | - | - | - |  |  |  |  | - | - | - | - | - | - |  | - | - | - | - | - | - | - |  | 0.5 | n/a | n/a | n/a | 0.3 | - | 0.2 | - | 0.3 | 0.3 | 0.3 | 0.3 | 0.4 | 0.5 | 0.3
Membrane Check
Last m (mg) |  | 0.849 | 0.848 | 0.852 | 0.851 | 0.849 | 0.85 | 0.849 | 0.855 | 0.855 | 0.821 | 0.859 | 0.855 | 0.857 | 0.86 | 0.855 | 0.871 | - | 0.85 | 0.875 | 0.854 | NA | NA | 0.858 |  | 0.786 | 0.78 | 0.781 | 0.781 | 0.782 | 0.78 | 0.782 | 0.783 | 0.783 | 0.783 | 0.785 | 0.785 | 0.79 | 0.788 | 0.786
ABS (mg) |  |  | 0.843 | 0.843 | 0.843 | 0.843 | 0.843 | 0.843 | 0.843 | 0.843 | 0.843 | 0.843 | 0.843 | 0.843 | 0.843 | 0.843 | 0.843 | 0.843 | 0.843 | 0.843 | 0.843 | 0.843 | 0.843 | 0.843 |  | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782
Difference (mg) |  |  | 0.005 | 0.009 | 0.008 | 0.006 | 0.007 | 0.006 | 0.012 | 0.012 | 0.022 | 0.016 | 0.012 | 0.014 | 0.017 | 0.012 | 0.028 | - | 0.007 | 0.032 | 0.011 |  |  | 0.015 |  | 0.004 | -0.002 | -0.001 | -0.001 | 0 | -0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0.002 | 0 | 0.004
[Last-org] / Org=% difference (<5%) |  |  | 0.5931% | 1.068% | 0.949% | 0.7117% | 0.8304% | 0.7117% | 1.423% | 1.423% | -2.61% | 1.898% | 1.423% | 1.661% | 2.017% | 1.423% | 3.321% | - | 0.8304% | 3.796% | 1.305% |  |  | 1.779% |  | 0% | -0.001% | 0.01% | 0.01% | 0% | 0.02% | 0% | 0% | 0% | 0% | 0% | 0% | 100% | 0% | 0%
Any status error flags showing? If so, list. | [25 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: 2022-08-25 to 2023-04-21
Is date/time correct within 60 seconds? | Y set | Y | Y | ADJ | Y | Y | Y | Y | ADJ | Y | ADJ | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y |  | yes | yes | yes | yes | corrected | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
Are all filter transport systems operational? | [25 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Are all tape reel cover knobs secure? | [25 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Are hourly data being collected? | [25 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Is the PM10 inlet head clean and secure? | OK | OK | OK | CLND | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y
Is the inlet head bug screen clean? | [25 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes, changed | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Is the water collection bottle clean? | [25 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes, changed | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Are o-rings in good condition? | [25 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Is there enough filter tape? (what % of the roll remains?) | 90% | 80% | 70% | 50% | 45% | 30% | 15% | 50% | 40% | 20% | 100% | 90% | 50% | 40% | 100% | 90% | 90% | 80% | 60% | 30%% | 20% | 90% | 70% | 90% |  | 50% | yes | yes | yes | yes | yes, changed | 75% | Around 50% | Just over 25% | 100% | 75%% | 60% | 45-50% | 25% | 100%
Does the filter tape show any pinholes or irregular PM deposits? | [25 empty] | no | no | no | no | no | no | pin holes | A few X's | pinhole, top R | small tiny hole started | No | no | no | no | no | bar: 2022-08-25 to 2023-04-21
Is the nozzle clean? | Y clnd | Y | CLND | Y | Y | CLND | Y | Y | CLND | Y | CLND | Y | CLND | Y | CLND | Y | Y | Y | CLND | Y | CLND | Y | Y | Y |  | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
Is the support vane clean? | Y clnd | Y | CLND | Y | Y | CLND | Y | Y | CLND | Y | CLND | Y | CLND | Y | CLND | Y | Y | Y | CLND | Y | CLND | Y | Y | Y |  | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
Is the capstan clean? | Y clnd | Y | CLND | Y | Y | CLND | Y | Y | CLND | Y | CLND | Y | CLND | Y | CLND | Y | Y | Y | CLND | Y | CLND | Y | Y | Y |  | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
Is the pinch roller clean? | Y clnd | Y | CLND | Y | Y | CLND | Y | Y | CLND | Y | CLND | Y | CLND | Y | CLND | Y | Y | Y | CLND | Y | CLND | Y | Y | Y |  | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
Is the pump muffler clean? | [25 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Is anything touching the (HOT) pump? | [25 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: 2022-08-25 to 2023-04-21
Is the temperature probe clean/tight? | [25 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Is the temperature probe signal connection tight? | [25 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Is the signal connections on the rear panel tight? | [25 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Is the AC air filter clean? | rotated | Y | Y | Y | - | - | - | - | Y | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - |  | yes | - | - | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | changed
Is the evap. coil (base of AC unit) clean? | clnd | Y | Y | Y | - | - | - | - | Y | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - |  | yes | - | - | - | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | cleaned
Is the AC comp. drain line clean? | [25 empty] | yes | - | - | - | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Is the AC temp. setting ~82F? | 78 |  | Y | - | - | - | - | - | ADJ | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - |  | yes | - | - | - | yes | yes | yes | yes | yes | yes | - | yes | yes | yes | yes
General Info
Is the site secure with no signs of damage? | [25 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Are all monitors in SAMPLE mode (no standby)? | [25 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Is communication system operational? | [25 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Is climate control working? Note current inside temp in Deg F | Y | Y | 77.4 | 76.55 | 22.81 | 23.21 | 15.36 | 17.23 | 13.8 | 25.85 | 17.02 | 16.26 | 16.4 | 16.8 | 25.69 | 22.47 | 21.69 | - | 24.63 | 16.17 | 20.1 | 25.3 | 36.38 | 25.46 |  | yes | - | yes | yes | 82F | yes | 82F | 83 | 83 | yes | yes | yes | yes | yes | yes
*Are there any unusual site noises? | [25 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: 2022-08-25 to 2023-04-21
*Are there any unusual site observations? | [25 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: 2022-08-25 to 2023-04-21
Is the site clean and manicured? | [25 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes, no fence | yes | yes | yes | yes | yes | tall grass | bar: 2022-08-25 to 2023-04-21
Are there any site safety issues? | [25 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: 2022-08-25 to 2023-04-21
*Are supplies and expendables needed or ordered? | [25 empty] | ok | ok | yes | yes | no | no | no | no | no | Long Q-tips | no | no | no | no | no | bar: 2022-08-25 to 2023-04-21
Operator Initials: | [25 empty] | JCG | jcg | gh | jcg | ds | jcg | ds | JCG | JCG | JCG | JCG | JCG | JCG | JCG | DS | bar: 2022-08-25 to 2023-04-21
Departure Time: | 15:30 | 10:30 | 11:40 | 11:15 | 11:50 | 12:20 | 10:30 | 11:00 | 11:10 | 2:15 | 12:35 | 10:30 | 12:25 | 12:20 | 1:00 | 12:20 | 11:30 | 9:35 | 12:30 | 11:00 | 10:30 | 12:20 | 11:45 | 11:25 |  |  | 1136 | 1201 | 1145 | 955 | 1155 | 15:15 | 1315 | 1340 | 1522 | 1405 | 1330 | 1210 | 1150 | 15:15
*If Yes call Dave Szatkowski to discuss (512)653-1737

## ST-N
KCBX ST-N Operator Checklist |  | 10 weeks
Arrival Date: (to insert new dates selet Column C. - Right click - insert 1 left) | 2022-08-25 | 2022-09-02 | 2022-09-02 | 2022-09-16 | 2022-09-21 | 2022-09-27 | 2022-10-07 | 2022-10-14 | 2022-10-19 | 2022-11-02 | 2022-11-14 | 2022-11-28 | 2022-12-13 | 2022-12-28 | 2023-01-05 | 2023-01-13 | 2023-01-20 | 2023-01-27 | 2023-02-06 | 2023-02-24 | 2023-03-10 | 2023-03-30 | 2023-04-14 | 2023-04-21 | 2015-10-28 | 2015-04-06 | 2015-04-14 | 2015-04-21 | 2015-05-06 | 2015-05-20 | 2015-06-03 | 2015-06-19 | 2015-07-06 | 2015-07-21 | 2015-08-04 | 2015-08-17 | 2015-09-03 | 2015-09-16 | 2015-10-01
Arrival Time: | 13:30 | 9:15 |  | 11:20 | 12:15 | 12:25 | 10:30 | 11:00 | 11:10 | 2:15 | 12:35 | 10:30 | 12:45 | 12:20 | 1:00 | 12:20 | 11:00 | 9:35 | 12:30 | 10:35 | 10:30 | 12:20 | 11:45 | 11:25 | 1111 |  |  | 1125 | 9:05 | 11:14 | 14:00 | 1245 | 1303 | 1427 | 1320 | 1255 | 1136 | 1058 | 13:40
BAM-1020 PM10 Monitor
Is it time for a monthly flow check? | [24 empty] | yes | yes | Yes | yes | yes | yes | yes | yes | yes | yes | yes | Yes | Yes | yes | yes | bar: 2022-08-25 to 2023-04-21
IF SO, what is the 'as found' sampler indicated flow rate (LPM)? | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 |  | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.7 | 16.7 | - | - | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.7 | 17 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7
what is indicated delta cal flow rate? | 16.7 | 16.71 | 16.82 | 16.62 | 16.73 | 16.63 |  | 16.62 | 16.78 | 16.68 | 16.71 | 16.64 | 16.59 | 16.64 | - | 16.72 | 16.66 | - | - | 16.95 | 16.4 | 17.05 | 16.57 | - | 16.52 | 16.79 | 16.66 | 16.71 | 16.67 | 16.76 | 16.3 | 16.4 | 16.2 | 16.5 | 16.6 | 16.51 | 16.8 | 16.65 | 16.83
what is the difference betwen cal and sampler flow (passing < 0.7 LPM) | 0 | 0.01 | 0.12 | 0.08 | 0.03 | 0.07 | 0 | 0.08 | 0.08 | 0.02 | 0.01 | 0.06 | 0.11 | 0.06 | - | 0.02 | 0.04 | - | - | 0.25 | 0.3 | 0.35 | 0.13 | - | 0.18 | 0.21 | 0.04 |  | 0.03 | 0.06 | 0.4 | 0.3 | 0.5 | 0.2 | 0.1 | 0.19 | 0.1 | 0.05 | 0.13
what is the difference betwen cal and design flow (passing < 0.8 LPM) | 0 | 0.01 | 0.12 | 0.08 | 0.03 | 0.07 | 0 | 0.08 | 0.08 | 0.02 | 0.01 | 0.06 | 0.11 | 0.06 | - | 0.02 | 0.04 | - | - | 0.25 | 0.3 | 0.35 | 0.13 | - | 0.28 | 0.09 | 0.04 |  | 0.03 | 0.06 | 0.4 | 0.3 | 0.6 | 0.3 | 0.2 | 0.29 | 0.2 | 0.15 | 0.13
IF ADJUSTMENTS ARE MADE, what is the 'as left' sampler rate (LPM)? | 16.7 | - | - | - | - | - |  | - | - | - | - | - | 16.6 | - | - | - | - | - | - | 16.9 | 16.4 | 17 | - | - | - | n/a | n/a | n/a | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.7
what is the 'as left' calibraton rate (LPM) | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 |  | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.7 | 16.7 | - | - | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.6 | 17 | 16.7 |  | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.8 | 16.7 | 16.6
Record the Bam's ambient temp | 23.3 | 26.3 | 27.9 | 26.7 | 26.6 | 16.2 |  | 11.8 | 8 | 22.5 | 6 | 8.1 | 6.7 | 6.1 | - | 0.2 | 0.9 | - | - | 7.4 | 2.4 | 6.7 | 25.1 | -
Record the Delta-cal ambient temp (passing = +/-2 deg C) | 24.7 | 27.8 | 28.5 | 27.7 | 27.6 | 11.3 |  | 9.2 | 9.6 | 22.3 | 5.7 | 8.8 | 6.4 | 6.8 | - | 0.8 | 1.8 | - | - | -2.8 | 1.2 | 6.3 | 24.8 | -
if adjustment is made, record the new sampler ambient temp | 24.7 | 27.8 | 28.5 | - | - | 11.3 |  | - | - | - | - | 8.8 | 6.4 | - | - | - | 1.8 | - | - | -2.8 | 1.2 | - | - | -
Record the Bam's BP | 745 | 749 | 747 | 750 | 744 | 749 |  | 739 | 747 | 752 | 754 | 746 | 741 | 742 | - | 751 | 743 | - | - | 759 | 745 | 751 | 242 | -
Record the Delta-cal BP(passing = +/- 10mmHg) | 747 | 749 | 747.9 | 750.5 | 743.5 | 749.5 |  | 738.5 | 747 | 752 | 754.5 | 746 | 741 | 741.4 | - | 750.2 | 748 | - | - | 761.1 | 745.2 | 752.5 | 242.5 | -
if adjustment is made, record the new sampler BP | 747 | - | - | - | - | - |  | - | - | - | - | - | - | - | - | - | 747.4 | - | - | 761 | - | - | - | -
Is it time for a monthly leak check? | [24 empty] | yes | no | yes | no | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
IF SO, what is the 'as found' leak check value (passing < 0.5 LPM)? | 0.1 | 0.2 | 0.3 | 0.2 | 0.2 | 0.3 |  | 0.3 | 0.3 | 0.3 | 0.4 | 0.3 | 0.3 | 0.4 | - | 0.4 | 0.3 | - | - | 0.4 | 0.4 | 0.5 | 0.4 | - | 0.4 | - | 0.4 | - | 0.3 | 0.3 | 2.8 | 0.1 | 1.3 | 0.4 | 0.3 | 0.3 | 0.4 | 0.3 | 0.4
IF FAILING, what is the 'as left' leak check value? | 0.1 | - | - | - | - | - |  | - | - | - | - | - | - | - | - | - |  | - | - | - | - | - | - | - | 0.5 | n/a | n/a | n/a | 0.3 | - | 0.2 | - | 0.3 | 0.3 | 0.3 | 0.3 | 0.4 | 0.5 | 0.3
Membrane Check
Last m (mg) | 0.764 | 0.768 | 0.768 | 0.772 | 0.769 | 0.771 | 0.77 | 0.769 | 0.767 | 0.77 | 0.77 | 0.769 | 0.762 | 0.767 | 0.77 | 0.77 | 0.767 | 0.769 | 0.846 | 0.849 | 0.846 | NA | 0.857 | 0.852 | 0.786 | 0.78 | 0.781 | 0.781 | 0.782 | 0.78 | 0.782 | 0.783 | 0.783 | 0.783 | 0.785 | 0.785 | 0.79 | 0.788 | 0.786
ABS (mg) | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 0.839 | 0.839 | 0.839 | 0.839 | 0.839 | 0.839 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782
Difference (mg) | 0.014 | 0.018 | 0.018 | 0.022 | 0.019 | 0.021 | 0.02 | 0.019 | 0.017 | 0.02 | 0.02 | 0.019 | 0.012 | 0.017 | 0.02 | 0.02 | 0.017 | 0.019 | 0.007 | 0.01 | 0.007 |  | 0.018 | 0.013 | 0.004 | -0.002 | -0.001 | -0.001 | 0 | -0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0.002 | 0 | 0.004
[Last-org] / Org=% difference (<5%) |  | 2.4% | 2.4% | 2.933% | 2.533% | 2.8% | 2.667% | 2.533% | 2.267% | 2.667% | 2.667% | 2.533% | 1.6% | 2.267% | 2.667% | 2.667% | 2.267% | 2.533% | 0.8343% | 1.192% | 0.8343% |  | 2.145% | 1.549% | 0% | -0.001% | 0.01% | 0.01% | 0% | 0.02% | 0% | 0% | 0% | 0% | 0% | 0% | 100% | 0% | 0%
Any status error flags showing? If so, list. | [24 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: 2022-08-25 to 2023-04-21
Is date/time correct within 60 seconds? | [24 empty] | yes | yes | yes | yes | corrected | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Are all filter transport systems operational? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Are all tape reel cover knobs secure? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Are hourly data being collected? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-25 to 2023-04-21
Is the PM10 inlet head clean and secure? | OK | OK | OK | CLND | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | RPLCD | Y | Y | Y | Y | Y | Y | Y | Y | Y
Is the inlet head bug screen clean? | OK | OK | OK | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | RPLCD | Y | Y | Y | Y | Y | Y | Y | Y | Y | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes, changed | yes | yes | yes
Is the water collection bottle clean? | OK | OK | OK | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | RPLCD | Y | Y | Y | Y | Y | Y | Y | Y | Y | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes, changed | yes | yes | yes
Are o-rings in good condition? | OK | OK | OK | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | Y | RPLCD | Y | Y | Y | Y | Y | Y | Y | Y | Y | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
Is there enough filter tape? (what % of the roll remains?) | 25% | 15% | 95% | 90% | 80% | 70%% | 60% | 50% | 30% | 20% | 90% | 80% | 50% | 30% | 20% | 100% | 90% | 90% | 70% | 30% | 20% | 95% | 70% | 60% | 50% | yes | yes | yes | yes | yes, changed | 75% | Around 50% | Just over 25% | 100% | 75%% | 60% | 45-50% | 25% | 100%
Does the filter tape show any pinholes or irregular PM deposits? | [24 empty] | no | no | no | no | no | no | pin holes | A few X's | pinhole, top R | small tiny hole started | No | no | no | no | no | bar: OK to Y
Is the nozzle clean? | Y clnd | Y | Y | Y | Y | CLND | Y | CLND | Y | CLND | Y | CLND | CLND | Y | Y | Y | CLND | Y | Y | Y | CLND | Y | Y | Y | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
Is the support vane clean? | Y clnd | Y | Y | Y | Y | CLND | Y | CLND | Y | CLND | Y | CLND | CLND | Y | Y | Y | CLND | Y | Y | Y | CLND | Y | Y | Y | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
Is the capstan clean? | Y clnd | Y | Y | Y | Y | CLND | Y | CLND | Y | CLND | Y | CLND | CLND | Y | Y | Y | CLND | Y | Y | Y | CLND | Y | Y | Y | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
Is the pinch roller clean? | Y clnd | Y | Y | Y | Y | CLND | Y | CLND | Y | CLND | Y | CLND | CLND | Y | Y | Y | CLND | Y | Y | Y | CLND | Y | Y | Y | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
Is the pump muffler clean? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: OK to Y
Is anything touching the (HOT) pump? | [24 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: OK to Y
Is the temperature probe clean/tight? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: OK to Y
Is the temperature probe signal connection tight? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: OK to Y
Is the signal connections on the rear panel tight? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: OK to Y
Is the AC air filter clean? | clnd | Y | Y | Y | - | Y | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | yes | - | - | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | changed
Is the evap. coil (base of AC unit) clean? | clnd | Y | Y | Y | - | CLND | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | yes | - | - | - | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | cleaned
Is the AC comp. drain line clean? | [24 empty] | yes | - | - | - | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: clnd to -
Is the AC temp. setting ~82F? | 80 |  | Lowered it to 80 | - | - | Y | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | yes | - | - | - | yes | yes | yes | yes | yes | yes | - | yes | yes | yes | yes
General Info
Is the site secure with no signs of damage? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: clnd to -
Are all monitors in SAMPLE mode (no standby)? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: clnd to -
Is communication system operational? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: clnd to -
Is climate control working? Note current inside temp in Deg F | Y | Y | 97.88 | 91.98 | 35.16 | 32.22 | 27.91 | 30.09 | 37.82 | 32.07 | 28.78 | - | 26.4 | 25.3 | 27.37 | 18.53 | 25.88 | 24.1 | 26.59 | 16.17 | 22.6 | 26.22 | 32.25 | 26.42 | yes | - | yes | yes | 82F | yes | 82F | 83 | 83 | yes | yes | yes | yes | yes | yes
*Are there any unusual site noises? | [24 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: clnd to -
*Are there any unusual site observations? | [24 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: clnd to -
Is the site clean and manicured? | [24 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes, no fence | yes | yes | yes | yes | yes | tall grass | bar: clnd to -
Are there any site safety issues? | [24 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no | bar: clnd to -
*Are supplies and expendables needed or ordered? | [24 empty] | ok | ok | yes | yes | no | no | no | no | no | Long Q-tips | no | no | no | no | no | bar: clnd to -
Operator Initials: | [24 empty] | JCG | jcg | gh | jcg | ds | jcg | ds | JCG | JCG | JCG | JCG | JCG | JCG | JCG | DS | bar: clnd to -
Departure Time: | 13:55 | 10:00 | 12:05 | 11:40 | 12:25 | 12:45 | 10:40 | 11:25 | 11:25 | 2:30 | 12:55 | 10:45 | 12:25 | 12:30 | 1:20 | 12:40 | 11:15 | 9:45 | 12:40 | 10:45 | 10:50 | 12:30 | 11:55 | 11:35 |  | 1136 | 1201 | 1145 | 955 | 1155 | 15:15 | 1315 | 1340 | 1522 | 1405 | 1330 | 1210 | 1150 | 15:15
*If Yes call Dave Szatkowski to discuss (512)653-1737

## ST-Met
South Terminal Met Station |  | 10 weeks
Arrival Date: (to insert new dates selet Column C. - Right click - insert 1 left) | 2022-08-26 | 2022-09-02 | 2022-09-08 | 2022-09-16 | 2022-09-21 | 2022-09-27 | 2022-10-07 | 2022-10-14 | 2022-10-19 | 2022-11-02 | 2022-11-14 | 2022-11-28 | 2022-12-13 | 2023-01-05 | 2023-01-13 | 2023-01-20 | 2023-01-27 | 2023-02-24 | 2023-03-30 | 2023-04-14 | 2015-10-28 | 2015-04-06 | 2015-04-14 | 2015-04-21 | 2015-05-06 | 2015-05-20 | 2015-06-03 | 2015-06-19 | 2015-07-06 | 2015-07-21 | 2015-08-04 | 2015-08-17 | 2015-09-03 | 2015-09-16 | 2015-10-01
Arrival Time: | 15:00 | 8:35 | 10:10 | 9:50 | 9:40 | 10:45 | 9:15 | 9:50 | 11:45 | 12:20 | 11:20 | 9:20 | 11:20 | 12:10 | 11:30 | 10:20 | 9:45 | 10:10 | 11:50 |  | 1111 |  |  | 1125 | 9:05 | 11:14 | 14:00 | 1245 | 1303 | 1427 | 1320 | 1255 | 1136 | 1058 | 13:40
[21 empty] | yes | yes | Yes | yes | yes | yes | yes | yes | yes | yes | yes | Yes | Yes | yes | yes
Is the wind speed sensor in good condition? | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | OK | DE-ICED | OK | OK | OK | 16.7 | 17 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7
Is the wind direction sensor in good condition? | [20 empty] | 16.52 | 16.79 | 16.66 | 16.71 | 16.67 | 16.76 | 16.3 | 16.4 | 16.2 | 16.5 | 16.6 | 16.51 | 16.8 | 16.65 | 16.83 | bar: 2022-08-26 to 2023-04-14
Is the temp aspirator in good condition? | [20 empty] | 0.18 | 0.21 | 0.04 |  | 0.03 | 0.06 | 0.4 | 0.3 | 0.5 | 0.2 | 0.1 | 0.19 | 0.1 | 0.05 | 0.13 | bar: 2022-08-26 to 2023-04-14
Is the aspirator fan working? | [20 empty] | 0.28 | 0.09 | 0.04 |  | 0.03 | 0.06 | 0.4 | 0.3 | 0.6 | 0.3 | 0.2 | 0.29 | 0.2 | 0.15 | 0.13 | bar: 2022-08-26 to 2023-04-14
Are the signal cables in good condition? | [20 empty] | - | n/a | n/a | n/a | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | - | 16.7 | bar: 2022-08-26 to 2023-04-14
Are the signal cables secured with tie wraps? | [20 empty] | 16.6 | 17 | 16.7 |  | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.7 | 16.8 | 16.7 | 16.6 | bar: 2022-08-26 to 2023-04-14
Does the wind speed appear accurate on the data logger? | bar: 2022-08-26 to 2023-04-14
Does the wind direction appear accurate on the data logger? | bar: 2022-08-26 to 2023-04-14
Does the temperature appear accurate on the data logger ? | bar: 2022-08-26 to 2023-04-14
Does the BP appear accurate on the data logger? | bar: 2022-08-26 to 2023-04-14
Record the 12V power supply reading | 13 | 13.32 | 13.32 | 13.32 | 13.32 | 13.32 | 13.33 | 13.23 | 13.25 | 13.37 | 13.23 | 13.37 | 13.3 | 13.19 | 13.32 | 13.31 | 13.3 | 13.2 | 13.3 | 13.2
Does the router have good signal strength? | bar: 2022-08-26 to 2023-04-14
Verify the comms antenna and cable are secure | [20 empty] | yes | no | yes | no | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | bar: 2022-08-26 to 2023-04-14
[21 empty] | 0.4 | - | 0.4 | - | 0.3 | 0.3 | 2.8 | 0.1 | 1.3 | 0.4 | 0.3 | 0.3 | 0.4 | 0.3 | 0.4
[21 empty] | 0.5 | n/a | n/a | n/a | 0.3 | - | 0.2 | - | 0.3 | 0.3 | 0.3 | 0.3 | 0.4 | 0.5 | 0.3
[21 empty] | 0.786 | 0.78 | 0.781 | 0.781 | 0.782 | 0.78 | 0.782 | 0.783 | 0.783 | 0.783 | 0.785 | 0.785 | 0.79 | 0.788 | 0.786
[21 empty] | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782 | 0.782
[21 empty] | 0.004 | -0.002 | -0.001 | -0.001 | 0 | -0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0.002 | 0 | 0.004
[21 empty] | 0% | -0.001% | 0.01% | 0.01% | 0% | 0.02% | 0% | 0% | 0% | 0% | 0% | 0% | 100% | 0% | 0%
[21 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no
[21 empty] | yes | yes | yes | yes | corrected | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
[21 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
[21 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
[21 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
[21 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes, changed | yes | yes | yes
[21 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes, changed | yes | yes | yes
[21 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
[21 empty] | 50% | yes | yes | yes | yes | yes, changed | 75% | Around 50% | Just over 25% | 100% | 75%% | 60% | 45-50% | 25% | 100%
[21 empty] | no | no | no | no | no | no | pin holes | A few X's | pinhole, top R | small tiny hole started | No | no | no | no | no
[21 empty] | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
[21 empty] | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
[21 empty] | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
[21 empty] | yes | yes | yes | yes | cleaned | yes | cleaned | yes | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned | cleaned
[21 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
[21 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no
[21 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
[21 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
[21 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
[21 empty] | yes | - | - | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | changed
[21 empty] | yes | - | - | - | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | cleaned
[21 empty] | yes | - | - | - | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
[21 empty] | yes | - | - | - | yes | yes | yes | yes | yes | yes | - | yes | yes | yes | yes
[21 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
[21 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
[21 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes | yes
[21 empty] | yes | - | yes | yes | 82F | yes | 82F | 83 | 83 | yes | yes | yes | yes | yes | yes
[21 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no
[21 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no
[21 empty] | yes | yes | yes | yes | yes | yes | yes | yes | yes, no fence | yes | yes | yes | yes | yes | tall grass
[21 empty] | no | no | no | no | no | no | no | no | no | no | no | no | no | no | no
[21 empty] | ok | ok | yes | yes | no | no | no | no | no | Long Q-tips | no | no | no | no | no
[21 empty] | JCG | jcg | gh | jcg | ds | jcg | ds | JCG | JCG | JCG | JCG | JCG | JCG | JCG | DS
[22 empty] | 1136 | 1201 | 1145 | 955 | 1155 | 15:15 | 1315 | 1340 | 1522 | 1405 | 1330 | 1210 | 1150 | 15:15
*If Yes call Dave Szatkowski to discuss (512)653-1737